## Sheet1
ErrorCode | Name | ShortDescription | Description
1536 | INFO_VERS_TERM | IPInfo::Version: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::Version attribute read.
1537 | INFO_VERS_CORRUPTED | IPInfo::Version: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::Version.
1538 | INFO_IDNUM_TERM | IPInfo::IDNumber: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IDNumber attribute read.
1539 | INFO_IDNUM_CORRUPTED | IPInfo::IDNumber: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IDNumber.
1540 | INFO_INITDATE_TERM | IPInfo::InitDate: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::InitDate attribute read.
1541 | INFO_INITDATE_CORRUPTED | IPInfo::InitDate: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::InitDate.
1542 | INFO_IPSIZE_TERM | IPInfo::IPSize: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPSize attribute read.
1543 | INFO_IPSIZE_CORRUPTED | IPInfo::IPSize: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPSize.
1544 | INFO_SCANSIZE_TERM | IPInfo::ScanSize: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::ScanSize attribute read.
1545 | INFO_SCANSIZE_CORRUPTED | IPInfo::ScanSize: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::ScanSize.
1546 | INFO_IPTYPE_TERM | IPInfo::IPType: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPType attribute read.
1547 | INFO_IPTYPE_CORRUPTED | IPInfo::IPType: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPType.
1548 | INFO_IPSENS_TERM | IPInfo::IPSensitivity: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPSensitivity attribute read.
1549 | INFO_IPSENS_CORRUPTED | IPInfo::IPSensitivity: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPSensitivity.
1550 | INFO_IPERA_TERM | IPInfo::IPErasibility: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPErasibility attribute read.
1551 | INFO_IPERA_CORRUPTED | IPInfo::IPErasibility: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPErasibility.
1552 | INFO_IPTHICK_TERM | IPInfo::IPThickness: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::IPThickness attribute read.
1553 | INFO_IPTHICK_CORRUPTED | IPInfo::IPThickness: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::IPThickness.
1554 | INFO_CHECKSUM_TERM | IPInfo::Checksum: unexpected attribute terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: No section terminator after IPInfo::Checksum attribute read.
1555 | INFO_CHECKSUM_CORRUPTED | IPInfo::Checksum: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::Checksum.
1556 | INFO_CHECKSUM_WRONG | IPInfo::Checksum: checksum error | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Checksum error for section IPInfo.
1557 | STATUS_VERS_TERM | IPStatus::Version: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPStatus::Version attribute read.
1558 | STATUS_VERS_NOTSUPPORTED | IPStatus::Version: not supported | The data on the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is written in a format that is not supported by the software installed on the digitizer. This is not a malfunction, but a compatibility issue between one or more cassette / IP units and the software installed on the digitizer. Detailed description: IPStatus::Version not supported.
1559 | STATUS_VERS_CORRUPTED | IPStatus::Version: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPStatus::Version.
1560 | STATUS_STATUS_TERM | IPStatus::Status: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPStatus::Status attribute read.
1561 | STATUS_STATUS_CORRUPTED | IPStatus::Status: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPStatus::Status.
1562 | STATUS_CYCLES_TERM | IPStatus::Cycles: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPStatus::Cycles attribute read.
1563 | STATUS_CYCLES_CORRUPTED | IPStatus::Cycles: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPStatus::Cycles.
1564 | STATUS_CHECKSUM_TERM | IPStatus::Checksum: unexpected attribute terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: No section terminator after IPStatus::Checksum attribute read.
1565 | STATUS_CHECKSUM_CORRUPTED | IPStatus::Checksum: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPStatus::Checksum.
1566 | STATUS_CHECKSUM_WRONG | IPStatus::Checksum: checksum error | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Checksum error for section IPStatus.
1567 | CTRL_VERS_TERM | IPCtrl::Version: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPCtrl::Version attribute read.
1568 | CTRL_VERS_NOTSUPPORTED | IPCtrl::Version: not supported | The data on the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is written in a format that is not supported by the software installed on the digitizer. This is not a malfunction, but a compatibility issue between one or more cassette / IP units and the software installed on the digitizer. Detailed description: IPCtrl::Version not supported.
1569 | CTRL_VERS_CORRUPTED | IPCtrl::Version: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPCtrl::Version.
1570 | CTRL_SENSITIVITY_TERM | IPCtrl::Sensitivity: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPCtrl::Sensitivity attribute read.
1571 | CTRL_SENSITIVITY_CORRUPTED | IPCtrl::Sensitivity: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPCtrl::Sensitivity.
1572 | CTRL_EXPDOSE_TERM | IPCtrl::ExposureDose: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPCtrl::ExposureDose attribute read.
1573 | CTRL_EXPDOSE_CORRUPTED | IPCtrl::ExposureDose: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPCtrl::ExposureDose.
1574 | CTRL_OPCODE_TERM | IPCtrl::OperationCode: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPCtrl::OpCode attribute read.
1575 | CTRL_OPCODE_CORRUPTED | IPCtrl::CoperationCode: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPCtrl::OpCode.
1576 | CTRL_CHECKSUM_TERM | IPCtrl::Checksum: unexpected attribute terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: No section terminator after IPCtrl::Checksum attribute read.
1577 | CTRL_CHECKSUM_CORRUPTED | IPCtrl::Checksum: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPCtrl::Checksum.
1578 | CTRL_CHECKSUM_WRONG | IPCtrl::Checksum: checksum error | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Checksum error for section IPCtrl.
1579 | IPPSDATA_VERS_TERM | IPPSData::Version: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::Version attribute read.
1580 | IPPSDATA_VERS_NOTSUPPORTED | IPPSData::Version: not supported | The data on the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is written in a format that is not supported by the software installed on the digitizer. This is not a malfunction, but a compatibility issue between one or more cassette / IP units and the software installed on the digitizer. Detailed description: IPPSData::Version not supported.
1581 | IPPSDATA_VERS_CORRUPTED | IPPSData::Version: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::Version.
1582 | IPPSDATA_MPPSINSTUID_TERM | IPPSData::MPPSInstanceUID: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::MPPSInstanceUID attribute read.
1583 | IPPSDATA_MPPSINSTUID_CORRUPTED | IPPSData::MPPSInstanceUID: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::MPPSInstanceUID.
1584 | IPPSDATA_SOPINSTUID_TERM | IPPSData::SOPInstanceUID: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::SOPInstanceUID attribute read.
1585 | IPPSDATA_SOPINSTUID_CORRUPTED | IPPSData::SOPInstanceUID: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::SOPInstanceUID.
1586 | IPPSDATA_PROCSTAT_TERM | IPPSData::ProcessingStation: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::ProcessingStation attribute read.
1587 | IPPSDATA_PROCSTAT_CORRUPTED | IPPSData::ProcessingStation: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::ProcessingStation.
1588 | IPPSDATA_STIMIDENTID_TERM | IPPSData::StimIdentID: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::StimIdentID attribute read.
1589 | IPPSDATA_STIMIDENTID_CORRUPTED | IPPSData::StimIdentID: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::StimIdentID.
1590 | IPPSDATA_EXPROOMID_TERM | IPPSData::ExposureRoomID: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPPSData::ExposureRoom attribute read.
1591 | IPPSDATA_EXPROOMID_CORRUPTED | IPPSData::ExposureRoomID: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::ExposureRoom.
1592 | IPPSDATA_CHECKSUM_TERM | IPPSData::Checksum: unexpected attribute terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: No section terminator after IPPSData::Checksum attribute read.
1593 | IPPSDATA_CHECKSUM_CORRUPTED | IPPSData::Checksum: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPPSData::Checksum.
1594 | IPPSDATA_CHECKSUM_WRONG | IPPSData::Checksum: checksum error | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Checksum error for section IPPSData.
1595 | INFO_VERSION_NOTSUPPORTED | IPInfo::Version: this version is not supported | The data on the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is written in a format that is not supported by the software installed on the digitizer. This is not a malfunction, but a compatibility issue between one or more cassette / IP units and the software installed on the digitizer. Detailed description: IPInfo::Version not supported.
1596 | INFO_APPTYPE_TERM | IPInfo::ApplicationType: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::AppType attribute read.
1597 | INFO_APPTYPE_CORRUPTED | IPInfo::ApplicationType: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::AppType.
1598 | IPCTRL_PATNAME_TERM | IPCtrl::PatientName: unexpected section terminator | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Section terminator after IPCtrl::PatientName attribute read.
1599 | IPCTRL_PATNAME_CORRUPTED | IPCtrl::PatientName: corrupted data stream | Identification data of the RF-tag mounted on the tray (in case of a flexible IP) or mounted on the IP (in case of a stiff IP) is corrupted. Please identify again (if identification by tablet is applicable), if this also fails, please exchange the RF-tag (or its carrier). Detailed description: Syntax error in attribute IPCtrl::PatientName.
1600 | INFO_OPTION_TERM | IPInfo::Option: unexpected section terminator | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Section terminator after IPInfo::Option attribute read.
1601 | INFO_OPTION_CORRUPTED | IPInfo::Option: corrupted data stream | Initialisation data of the RF-tag on the tray (for flexible IPs) or on the IP (stiff IPs) is corrupted. The RF-tag (or its carrier) has to be exchanged or initialized again, if possible. Detailed description: Syntax error in attribute IPInfo::Option.
4100 | ERASE_BUTTON_PRESSED_DURING_SCAN | erase button was pressed during scan cycle | Meaning: The erase button was pressed during the scan cycle. This is not allowed! Reason: The erase button was pressed too late: To start an erasure cycle, the erase button must be pressed before the cassette is inserted! Cure: Now identify on the processing station, then the scan cycle continues and the IP will be scanned and erased.
4101 | IDENTIFICATION_TIMEOUT | user did not identify at the processing station in time | Meaning: The digitizer is waiting for the identification data and the timeout (60 sec.) exceeded. Reason: User did not identify at the processing station in time. Cure: Identify at the processing station. If this is not working, check ethernet connection to the processing station and the network configuration.
4102 | IDENTIFICATION_FAILED_BUSY | The identification at the processing station failed because a previous ID request is still busy | Meaning: The identification at the processing station failed because a previous ID request is still busy and the digitizer tried several times to get the identification data from the processing station. Reason: Processing station hang-up, configuration error or software error. Cure: 1. Restart processing station and digitizer. 2. Check that processing station is correctly configured for this digitizer; check configuration of examination studies; if everything is correctly configured, create a backup to save the log files and inform both R&amp;D teams of processing station and digitizer.
4103 | IDENTIFICATION_FAILED | The identification at the processing station failed | Meaning: The identiication at the processing station failed. Reason: Processing station configuration error or software error. Cure: Check that the processing station is correctly configured for this digitizer; check configuration of examination studies. If everything is correctly configured, create a backup to save the logfiles and inform both R&amp;D teams of processing station and digitizer.
4104 | PREPARE_IPC_FAILED | Prepare image processing failed | Meaning:A necessary shading calibration file is missing or does not fit to the IP type or format. Reason: Missing shading calibration, because PMT with light collector or optic module was changed or calibration file is corrupt. Cure: Execute shading calibration
4105 | CONFIRM_CASSETTE_WITH_IP_FOUND_AT_START_UP | Start-Up: Cassette with IP found, no IP is in slowscan or transport unit | Meaning: Cassette with an IP found during start-up. Reason: Digitizer was switched off during cycle. Cure: Confirm message, remove cassette; check if image was sent to the processing station; if not, scan cassette again; erase cassette again in any case.
4108 | WRONG_IHC_PROTVER | Protocol version of IHC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of IHC (IP-HandlingControl) is wrong. Cure: Install new software; inform R&amp;D team.
4111 | WRONG_ICC_PROTVER | Protocol version of ICC is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of ICC (ImageCaptureControl) is wrong. Cure: Install new software; inform R&amp;D team.
4130 | WRONG_UID_PROTVER | Protocol version of UID is wrong | Meaning: Incompatible SW components. Reason: SW error, protocol version of UID (UserInterfaceDispatcher) is wrong. Cure: Install new software; inform R&amp;D team.
4155 | NX_NOT_AVAILABLE | processing station is not available | Meaning: The image processing application resp. the processing station is not available anymore. Reason: Image processing application is stopped at the processing station or a network problem occurred. Cure: Check if image processing application is running properly on the processing station, check ethernet connection to the processing station and the network configuration.
4167 | CAN_ADAPTER_ERROR | fatal error in CAN adapter layer | Meaning: Fatal error in CAN adapter layer Reason: CAN adapter layer cannot be initialized due to SW or HW error. Cure: Check supply voltages and CAN connections; update SW; if this doesn't help replace PMI board.
4168 | LOGGING_ADAPTER_ERROR | fatal error in logging adapter layer | Meaning: Fatal error in logging adapter layer Reason: Logging adapter layer cannot be initialized due to SW or HW error. Cure: Check supply voltages and CAN connections; update SW; if this doesn't help replace PMI board.
4170 | CAN_VCC_FAILURE | CAN Vcc failure | Meaning: CAN voltage error. Reason: - 'Cable PMI Board to Power Board' broken - 'Power Board' defective - 'PMI Board' defective Cure: 1. Restart the Digitizer and check, if the problem persists 2. Replace 'Cable PMI Board to Power Board' 3. Replace 'Power Board' 4. Replace 'PMI Board'
4173 | ANALOG_VCC_FAILURE | Analog Vcc failure | Meaning: The check of the +8 V supply voltage for A/D conversion on the PMI Board failed. Reason: - 'Cable PMI Board to Power Board' broken - 'Power Board' defective - 'PMI Board' defective Cure: 1. Restart the Digitizer and check, if the problem persists 2. Replace 'Cable PMI Board to Power Board' 3. Replace 'Power Board' 4. Replace 'PMI Board'
4174 | OPTICS_VCC_FAILURE | optics +12V or -12V failure | Meaning: The check of the supply voltage for the optic module on the PMII Board failed. Reason: - 'Cable PMI Board to Power Board' broken - 'Power Board' defective - 'PMI Board' defective Cure: 1. Restart the Digitizer and check, if the problem persists 2. Replace 'Cable PMI Board to Power Board' 3. Replace 'Power Board' 4. Replace 'PMI Board'
4198 | DIGITIZER_START_UP | digitizer is starting-up | Meaning: Digitizer is starting-up. Reason: Digitizer was switched on and is now starting-up. Cure: Wait unitil the start-up is finished
4199 | TIMEOUT_SCAN_END_ICC | timeout for 'scan' response from ICC after slowscan stopped | Meaning: The scanning unit didn't finish the scan process although the slowscan motor has already stopped. Reason: ICC (ImageCaptureControl) is still waiting for end of the scan, i.e. the required number of pixels have not been stored yet. Cure: Problem with the scanning unit (e.g. BeginOfLine/EndOfLine detection problem) 1. Check if acrylic light collector and PMT are mounted properly 2. Check supply voltages 3. Check cable connections between PMI board and PMT module; if OK, replace PMI 4. Replace PMT and light collector 5. Replace optic module
4200 | IP_FORMAT_NOT_VALID | Not supported IP format detected | Meaning: The IP format is not supported in this digitizer. Reason: Wrong IP format. Cure: Switch off the digitizer, remove image plate and cassette manually. Then press the cassette release button. Refer to the detailed description in the user manual. Use only cassettes with supported IP formats.
4201 | IP_APPLICATION_NOT_VALID | Not supported IP application detected | Meaning: The IP application is not supported in this digitizer. Reason: Wrong IP application. Cure: Press the cassette release button and remove cassette. Refer to the detailed description in the user manual. Use only cassettes with supported IP applications.
4202 | DIGITIZER_WAKE_UP | Digitizer wake-up, after houers not used | Meaning: Digitizer wake-up after Galvo was in sleep mode. Reason: Digitizer was not used for a few houers. Cure: Wait unitil the wake-up is finished
4219 | ICC_DID_NOT_DETECT_BOS | Reason: Wrong IP format. | Meaning: The scanning unit didn't start the scan process although the slowscan motor has finished the scan already. Reason: ICC (ImageCaptureControl) is still waiting for begin of the scan, i.e. no pixels have been stored yet. Cure: Problem with the scanning unit (e.g. BeginOfLine/EndOfLine detection problem) 1. Check supply voltages 2. Check cable connections between PMI board and optic module; 3. Replace PMI board 4. Replace Optic Module.
4220 | IPSTATUS_NOT_EXPOSED | Cure: Use cassette with supported IP format | Meaning: The IP cannot be scanned because the IP status is wrong. Reason: Digitizer rejects the scan because the processing station did not set the IP status to 'exposed'. Normally the IP status is set to 'exposed' by pressing the 'ID' button on the processing station. This is a software error in the processing station software. Cure: 1. Check if digitizer is configured correctly on the processing station 2. Check processing station software installation; if OK, inform R&amp;D team.
4221 | PMM_5V | Check PMM 5V failed | Meaning: The check of the +5V supply voltage for the PMT on the PMII Board failed. Reason: - 'Cable PMI Board to Power Board' broken - 'Power Board' defective - 'PMI Board' defective Cure: 1. Restart the Digitizer and check, if the problem persists 2. Replace 'Cable PMI Board to Power Board' 3. Replace 'Power Board' 4. Replace 'PMI Board'
4227 | PMM_ANALOG_POWER_FAILURE | PMM analog power +20V or -10V failure | Meaning: The check of the PMM analog power +20V and -10V on the PMI Board failed. Reason: - 'Cable PMI Board to Power Board' broken - 'Power Board' defective - 'PMI Board' defective Cure: 1. Restart the Digitizer and check, if the problem persists 2. Replace 'Cable PMI Board to Power Board' 3. Replace 'Power Board' 4. Replace 'PMI Board'
4229 | INITIALIZE_CAN | Initialize CAN BUS failed | CAN Bus initialisation failed. Meaning: Initialisation of the CAN bus software driver failed; Reason 1: CAN software driver failure; Cure 1: Install new digitizer software. Reason 2: CAN hardware on 'PMI Board' defective. Cure 2: Replace 'PMI Board'.
4230 | DIGITIZER_SW_VERSION_MISMATCH | The software versions of the internal digitizer modules do not match. | Meaning: Software versions loaded onto internal digitizer modules are not matching. Reason: Incomplete software installation, installation failed or hardware replacement (digitizer boards). Cure: Install digitizer software again.
4231 | INVALID_BUFFER_FROM_SPI | spooler returned invalid buffer | Meaning: Software component 'spooler' returned invalid buffer (= NULL). Reason: Fatal software error. Cure: Installl digitizer software again.
4232 | IMAGE_QUEUE_FULL | spooler input queue is full | Meaning: Image queue of software component 'spooler' is full. Reason:The spooler's capacity is exhausted, no image can be scanned at the moment. Cure: Wait until images are transmitted to the processing station. Be sure that the processing station is connected, configured correctly and working properly.
4233 | UNKNOWN_DESTINATION | destination is not known | Meaning: The destination is not known. Reason: The configured destination (= name of the processing station) is not known by the digitizer. Cure: Wrong digitizer or destination setting. Connect to service interface via browser and check the digitizer and destination settings (network addresses and names, AE titles, ports).
4234 | SNPRODUCT_EQ_SNUSB_NE_SNPMI | New PMI board detected. | Meaning: The device serial number stored on the PMI board for synchronization with USB flash drive does not correspond to the device serial number store in the device E-Label. Reason: A new PMI board has been installed for the digitizer or data on flash memory of PMI board are destructed. Cure: 1. If PMI board has been exchanged: configure the digitizer 2.If PMI board has not been exchanged: exchange PMI board
4235 | SNPRODUCT_NE_SNUSB_EQ_SNPMI | The USB flash drive is initialized for another device | Meaning: The device serial number stored on the USB flash drive does not correspond to the device serial number stored in the device E-Label of the digitizer. Reason: The user has inserted a USB flash drive from another digitizer. Cure: Follow the displayed instructions.
4236 | SNPRODUCT_NE_SNUSB_NE_SNPMI | The USB flash drive is initialized for another device, a new PMI has been detected. | Meaning: The device serial number stored on the USB flash drive does not correspond to the device serial number store in the Device E-Label of the digitizer. Additionally the device serial number on PMI does not correspond to the device serial number on the Device e_Label of the digitizer. Reason: The User has inserted a USB flash drive from another digitizer and the aPMI board has been exchanged. Cure: 1. Insert correct USB flash drive and restart the device to copy configuration data from USB flash drive to PMI board. 2. Configure PMI board for use for this digitizer, restart the digitizer and follow the instructions on processing station to use the insterted stick for the new configured PMI board.
4237 | NO_USB_INSERTED | No USB stick inserted. | Meaning: No USB flash drive has been detected by the device Reason: No USB flash drive is present or not completely inserted Cure: Insert USB flash drive correctly and restart device.
4238 | NO_USBID_SNPRODUCT_EQ_SNPMI | USB flash drive contains no ID data, insert new stick or initialize. | Meaning: No valid ID Data found on USB flash drive. Reason: USB flash drive does not contain valid ID data Cure: Follow the displayed instructions: Insert correct USB flash drive or initialize inserted USB flash drive (all data on USB flash drive will be destroyed!)
4239 | NO_USBID_SNPRODUCT_NE_SNPMI | USB flash drive contains no ID data, serial number on PMI does not correspond to serial number on device E-Label. | Meaning: No valid ID Data found on USB flash drive and no valid ID data on PMI board. Reason:New USB flash drive inserted and PMI board exchanged. Cure: 1. Insert correct USB flash drive and restart the device to copy configuration data from USB flash drive to PMI board. 2. Configure PMI board for use for this digitizer, restart the digitizer and follow the instructions on processing station to use the insterted stick for the new configured PMI board.
4240 | DATA_COPIED_RESET | Data have been copied successfully, please reset the device | Meaning: Copy action as indicated by previous step in displayed instructions was successfull. When pressing confirm the digitizer is restarted and should be ready to work.
4241 | USE_STICK | Do you wish to use inserted stick for actual device? All data will be lost | Meaning: No valid ID Data found on USB flash drive, the user is asked to confirm that he wants to use the inserted USB flash drive Reason: USB flash drive does not contain valid ID data Cure: Follow the displayed instructions.
4246 | SNPRODUCT_NOTINIT | Device E-Label not initialized yet | Meaning: The Device E-Label contains no valid data. Reason: The device eLabel is empty, was formatted or contains invalid data. Cure: Format the Device E-Label via ESF, then write valid data to the Device E-Label via ESF.
4248 | COULD_NOT_WRITE_SNELABEL_TO_PMI | Error writing device serial number to PMI Board | Meaning: Writing device serial number to PMI Board failed. Reason: Flash memory on PMI Board is full or defective. Cure: Install the SW with formatting of the flash memory on PMI Board. Restore previously saved data.
4249 | CANNOT_ACCESS_DEVICE_ELABEL | Error accessing Device E-Label | Meaning: The device E-Label is not detected. Reason: The device E-Label is not connected. Cure: Check cabling connection between PMI board and device E-Label.
4250 | CANNOT_ACCESS_OPTICS_ELABEL | Error accessing E-Label of Optic Module. | Meaning: The E-Label of the Optic Module is not detected. Reason: The E-Label of the Optic Module is not connected or defective. Cure: 1. Check cabling connection between PMI board and optic module. 2. Replace PMI board. 3. Replace optic module
4251 | IP_DATE_EXPIRE_SOON | IP date expiry date is reached soon | Meaning: The image plate expires soon. Reason: The expiry date of the image plate is nearly reached. Cure: Replace the image plate soon.
4252 | IP_DATE_EXPIRED | IP date expired | Meaning: The image plate expired, the system performance is worse than specified. Reason: The image plate expired. Cure: Replace the image plate.
4254 | CONFIRM_SPEEDFACTOR_OUT_OF_RANGE | Found Slow Scan speed factor out of range | Meaning: The scan length might not be within the required tolerance range. Reason: There may be distortions in the scanned image. Cure: Check the image plate for damages.
4255 | SPEEDFACTOR_FAR_OUT_OF_RANGE | Found Slow Scan speed factor far out of range | Meaning: The scan length is far out of the allowed tolerance range. Reason: Image plate damaged or mechanical problem like slippage in slow scan drive. Cure: 1. Check the image plate for damages. 2. Check belt tension of slow scan drum drive. 3. Check slow scan mechanics. 4. Check cable to IP detection sensor. 5. Replace IP detection sensor.
4256 | SNPRODUCT_EQ_SNUSB_SNPMI_CORRUPTED | Device serial number on PMI board not readable. | Meaning: Device serial number on PMI board could not be read Reason: The file d:/ProductInfo.xml is corrupted. Cure: Reinstall software with option format.
4257 | OUT_OF_MEMORY_READING_SN_PMI | Error reading device serial number on PMI board | Meaning: Device serial number on PMI board could not be read Reason: Software error Cure: Please reinstall software with option format.
4258 | OUT_OF_MEMORY_READING_SN_USB | Error reading device serial number on USB flash drive | Meaning: Device serial number on USB flash drvie could not be read Reason: Software error Cure: Please reinstall software with option format.
4259 | OUT_OF_MEMORY_READING_SN_ELABEL | Error reading serial number on Device E-Label | Meaning: Serial number on Device E-Label could not be read Reason: Software error Cure: Please reinstall software with option format.
4260 | NO_USBID_NO_SNPRODUCT_PMI | No device serial number on USB flash drvie and no device serial number on PMI board found | Meaning: No device serial number on USB flash drive and no device serial number on PMI board found Reason: 1. An empty USB flash drive has been inserted or USB flash drive is not readable and a new PMI board has been installed in the device. 2. An empty USB flash drive has been inserted or USB flash drive is not readable and data on PMI board have been destroyed. Cure: 1. Use the correct USB flash drive and press confirm to use configuration data from USB flash drive. 2. If you already are using the correct USB flash drive and this message is displayed a second time, reinstall the software with option format.
4261 | SNPRODUCT_EQ_SNUSB_NO_SNPRODUCT_PMI | Serial number on device E-Label identical to serial number on USB flash drive, no device serial number on PMI board found | Meaning: Serial number on Device E-Label identical to serial number on USB flash drive, no device serial number on PMI board found Reason: 1. A new PMI board has been installed in the device. 2. Data on PMI board have been destroyed. Cure: Please reinstall software with option format.
4262 | SNPRODUCT_NE_SNUSB_NO_SNPRODUCT_PMI | Serial number on device E-Label is different from device serial number on USB flash drive, no device serial number on PMI board found | Meaning: Serial number on Device E-Label is different from device serial number on USB flash drive, no device serial number on PMI board found Reason: 1. An USB flash drive from another device has been inserted and a new PMI board has been installed in the device. 2. An USB flash drive from another device has been inserted and data on PMI board have been destroyed. Cure: 1. Use the correct USB flash drive and press confirm to use configuration data from USB flash drive. 2. If you already are using the correct USB flash drive and this message is displayed a second time, reinstall the software with option format.
4263 | SNPRODUCT_NE_SNUSB_SNPMI_CORRUPTED | Device serial number on PMI board not readable. | Meaning: Device serial number on PMI board could not be read Reason: The file d:/ProductInfo.xml is corrupted. Cure: Please reinstall software with option format.
4264 | NO_USBID_SNPMI_CORRUPTED | Device serial number on PMI board not readable. | Meaning: Device serial number on PMI board could not be read Reason: The file d:/ProductInfo.xml is corrupted. Cure: Please reinstall software with option format.
4265 | USE_USB_CONFIG_FOR_DEVICE | Use configuration data from USB flash drive for this digitizer. | Meaning: A USB flash drive has been inserted that was initialized for another digitizer. The user is requested to decide whether to use the configuration from USB flash drive for this digitizer or cancel operation Cure: Follow the displayed instructions.
4266 | IDENTIFICATION_FAILED_WRONG_APPTYPE | Workstation sends wrong application type. | Meaning: The identiication at the processing station failed. Reason: Digitizer received the wrong application type (scan resolution) from the processing station. This can be a configuration error or a software error. Cure: Check that the processing station is correctly configured for this digitizer; check configuration of examination studies. If everything is correctly configured, create a backup to save the logfiles and inform both R&amp;D teams of processing station and digitizer.
4267 | IDENTIFICATION_FAILED_UNKNOWN_PS | Workstation sends unknown processing station name. | Meaning: The identiication at the processing station failed. Reason: Digitizer received an unknown processing station name from the processing station. This can be a configuration error or a software error. Cure: Check that the processing station is correctly configured for this digitizer. If everything is correctly configured, create a backup to save the logfiles and inform both R&amp;D teams of processing station and digitizer.
4269 | IDENTIFICATION_FAILED_CHECKSUM | The identification at the processing station failed due to a checksum error. | Meaning: The identiication at the processing station failed due to a checksum error. Reason: Checksum of received identification data is wrong or corrupt. Cure: Restart PC and digitizer and repeat the cycle. If error remains, create a backup to save the logfiles and inform both R&amp;D teams of processing station and digitizer.
4270 | IP_TYPE_NOT_VALID | Unsupported IP type detected. | Meaning: The IP type is not supported in this digitizer. Reason: Wrong IP type. Cure: Press the cassette release button and remove cassette. Refer to the detailed description in the user manual. Use only cassettes with supported IP types.
4271 | NO_USB_LICENSE_DRIVE_INSERTED | No USB license drive inserted. | Meaning: This is a licensed digitizer, but the USB license drive was not detected by the device. Reason: The USB license drive is not present or not inserted correctly. Cure: Insert USB license drive correctly and restart device.
4272 | UKNOWN_OPTICS_PROTECTION_STATE | The licensing state of the optics module is not known. | Meaning: The licensing state of the optics module is not known. Reason: SW error, licensing state is not determined correctly. Cure: Install new software. Inform your local support organization.
4273 | UKNOWN_PMM_PROTECTION_STATE | The licensing state of the PMT with light collector is not known. | Meaning: The licensing state of the PMT with light collector is not known. Reason: SW error, licensing state is not determined correctly. Cure: Install new software. Inform your local support organization.
4274 | UKNOWN_DEVICE_PROTECTION_STATE | The licensing state of the device is not known. | Meaning: The licensing state of the device is not known. Reason: SW error, licensing state is not determined correctly. Cure: Install new software. Inform your local support organization.
4275 | PERIODIC_LICENSE_EXPIRES_SOON | The license on USB license drive expires soon. | Meaning: The periodic license on the USB license drive expires soon. Reason: The expiry date of periodic license is nearly reached. Cure: Visit 'licensing.healthcare.agfa.net' and follow the instructions to update your license.
4276 | PERIODIC_LICENSE_EXPIRED | The license on USB license drive has expired. | Meaning: The periodic license on the USB license drive has expired. Reason: Periodic license expired. Cure: Update the periodic license.
4277 | FATAL_ERR_WITHOUT_CODE | A fatal error occured, but the error code is not set. | Meaning: A fatal error occured, but the error code is not set (i.e. it is 0) Reason: SW error, error code is not set. Cure: Install new software. Inform your local support organization.
4278 | PROTECTED_OPTICS_IN_UNPROTECTED_DEVICE | A protected optics module is installed in an unprotected device. | Meaning: An optics module, which is protected by a periodic license, is installed in an unprotected device. This is not allowed. Reason: A protected optics module was installed in an unprotected device. Cure: Check if the optics module is marked as a protected module. Replace protected optics module by an unprotected module (spare part).
4279 | PROTECTED_PMM_IN_UNPROTECTED_DEVICE | A protected PMT with light collector is installed in an unprotected device. | Meaning: A PMT with light collector, which is protected by a periodic license, is installed in an unprotected device. This is not allowed. Reason: A protected PMM was installed in an unprotected device. Cure: Check if the PMM is marked as a protected module. Replace protected PMM by an unprotected module (spare part).
4280 | WRONG_PROTECTED_OPTICS_IN_PROTECTED_DEVICE | A protected optics module from another licensed digitizer is installed in this protected device. | Meaning: An optics module, which is protected by a periodic license for another digitizer, is installed in this licensed device. This is not allowed. Reason: A protected optics module from another licensed digitizer was installed in this licensed device. Cure: Check if the optics module is marked as a protected module and in which digitizer it was installed before. Replace protected optics module by an unprotected module (spare part).
4281 | WRONG_PROTECTED_PMM_IN_PROTECTED_DEVICE | A protected PMT with light collector from another protected digitizer is installed in this protected device. | Meaning: A PMT with light collector module, which is protected by a periodic license for another digitizer, is installed in this protected device. This is not allowed. Reason: A protected PMT with light collector from another digitizer was installed in this protected device. Cure: Check if the PMT with light collector is marked as a protected module and in which digitizer it was installed before. Replace protected optics module by an unprotected module (spare part).
4282 | INVALID_PRODUCT_SERIAL_NUMBER | The determined product serial number is invalid. | Meaning: The determined product serial number is 0 and therewith invalid. Reason: SW error, product serial number is not determined correctly. Cure: Install new software. Inform your local support organization.
4283 | ALLOWED_DIGITIZER_TIME_NULLPTR | The SW method which determines the allowed digitizer time got a null pointer. | Meaning: The SW method which determines the allowed digitizer time got a null pointer. Reason: SW error, pointer must not be null. Cure: Install new software. Inform your local support organization.
4284 | DIFF_DONGLE_SYSTEM_CERTIFIED_TIME | The difference between system time and certified time on the USB license drive is too big. | Meaning: Certified time on USB license drive is older as the Digitizer system time. Digitizer is operational. Reason: 1) The Digitizer date and time is locked and can not be modified. The delay of the certified time on USB license drive to the Digitizer system time is out of limit. Cure: 1) First align the system time of the USB license drive (update time certificates): Visit 'licensing.healthcare.agfa.net' and update the reference time of the USB license drive. 2) Set the system date and time of the Digitizer accordingly in the Service User Interface, refer to the Service Manual.
4285 | DIGITIZER_TIME_NOT_ALLOWED | Setting the digitizer date and time to the entered value is not allowed. | Meaning: The digitizer date and time cannot be set to the entered value, because the digitizer is protected by an USB license drive which does not allow to set this digitizer date and time. Reason: The entered digitizer date and time is not allowed by the USB license drive. Cure: Update the periodic license and set the digitizer time correctly.
4286 | CERTIFIED_TIME_NEWER_THAN_DONGLE_SYSTEM_TIME | The certified time is newer than the system time on the USB license drive. | Meaning: The certified time is newer than the system time on the USB license drive. This should never occur. Reason: The time system of the USB license drive is disturbed. Cure: 1. Visit 'licensing.healthcare.agfa.net' and update the reference time of the USB license drive. 2. Set the system date and time of the Digitizer accordingly in the Service User Interface, refer to the Service Manual.
4298 | CLEANUP_FINISHED_WITH_CASSETTE_AND_IP_DETECTED | CleanUp terminates as expected at 'Cassette and IP detected' | The CleanUp procedure finishes successful with the found IP back into the cassette, ready for removal.
4299 | CLEANUP_MSG_AT_SITUATION_NO_CASS_AND_IP_DETECTED | CleanUp message at 'No-Cassette present and IP detected' | The CleanUp procedure found the situation 'No-Cassette present and IP inside the digitizer detected' and indicates the status with a message.
4300 | CLEANUP_FIN_AT_SITUATION_NO_CASS_AND_IP_DETECTED | CleanUp terminates as expected at 'No-Cassette present and IP detected' | The CleanUp procedure finishes successful with the found IP back into the inserted cassette, ready for removal.
4301 | CLEANUP_FIN_AT_SITUATION_CASS_AND_NO_IP_FOUND | CleanUp terminates as expected at 'Cassette present and No-IP found' | The CleanUp procedure finishes successful with the inserted cassette found empty, ready for removal.
4302 | CLEANUP_FINISHED_WITH_FAILURE | CleanUp terminates not successful | The CleanUp procedure reached an unexpected failure situation during a performed step. Root Cause stored in Infocounter Error History.
4303 | CLEANUP_FINISHED_TESTMODE_FAILURE | CleanUp terminates successful, testmode failure | The CleanUp procedure finished, but an unexpected failure situation is reached. Testmode is active. Startup blocked.
4304 | DIGITIZER_START_ADJUSTMENT | Digitizer is adjusting during start-up | Meaning: Digitizer is adjusting its optics. Reason: Digitizer adjust necessary. Cure: Wait a few minutes, the message will disappear automatically.
4305 | DIGITIZER_ADJUSTMENT | digitizer is adjusting | Meaning: Digitizer is adjusting its optics. Reason: Digitizer adjust necessary. Cure: Wait a few minutes, the message will disappear automatically.
4306 | ERASE_FINISHED_WITH_FAILURE | Erase terminates not successful | The Erasure procedure reached an unexpected failure situation during a performed step. Root Cause stored in Infocounter Error History.
4307 | RETURN_IP_FINISHED_WITH_FAILURE | Return IP terminates not successful | The ReturnIP procedure reached an unexpected failure situation during a performed step. Root Cause stored in Infocounter Error History.
4308 | TRANSPORT_IP_TO_SCANNER_FAILURE | Transportation of the IP into the scanner was not successful | The transport IP to scanner procedure reached an unexpected failure situation during a performed step. Additional info to the root cause is stored in Infocounter Error History. Confirm the user text message to proceed.
4309 | TRANSP_IP_TO_SCANNER_PREP_WITH_CLEANUP_FAILURE | Transport IP into the scanner problem. Prepare digitizer with CleanUp failed. | Transport IP to scanner procedure reached an unexpected failure situation. Prepare Digitzer again with CleanUp failed. Root Cause stored in Infocounter Error History.
4310 | TRANSP_IP_TO_BEGIN_OF_ERA_FAILURE_PREP_HAS_NO_IP | Transport IP to the begin of erasure problem. Prepare digitizer with CleanUp has no IP found. | Transport IP to begin of erasure at erasure mode reached an unexpected failure situation. Prepare Digitzer again with CleanUp found no IP inside. Root Cause stored in Infocounter Error History.
4311 | TRANSP_IP_TO_BEGIN_OF_ERA_FAILURE_PREP_HAS_IP | Transport IP to the begin of erasure problem. Prepare digitizer with CleanUp has an IP found. | Transport IP to begin of erasure at erasure mode reached an unexpected failure situation. Prepare Digitzer again with CleanUp found an IP inside. Root Cause stored in Infocounter Error History.
4312 | TRANSPORT_IP_TO_SCANNER_FAILURE_PREP_HAS_IP | Transportation of the IP into the scanner was not successful. Prepare digitizer with CleanUp has an IP found. | Transport IP to scanner procedure reached an unexpected failure situation. Prepare Digitzer again with CleanUp found an IP inside. Root Cause stored in Infocounter Error History.
4313 | OPTICS_BOARD_CONNECTOR | Optics board not detected | Meaning: Optic board not detected. Reason: Cable broken/not connected or optic board defective. Cure: 1) Check optic module cabling. 2) Replace optic module cabling, if error persists. 3) Replace optic module, if error persists.
4314 | PMM_BOARD_CONNECTOR | PMM board not detected | Meaning: PMM board not detected. Reason: PMM is not connected to PMI board. Cure: 1) Check PMM cabling. 2) Replace PMM cabling, if error persists. 3) Replace PMM, if error persists.
4315 | CANNOT_OPEN_ERROR_INFO_FILE | Cannot open error information file | Meaning: Could not open error information file while trying to display an error and ErrorViewer is not accessible. Reason: 1) USB flash memory not inserted. 2) USB flash memory has no free memory left 3) USB flash memory is write protected or does not allow write access for any other reason Cure: 1) Insert a USB flash memory with 10 kBytes free memory left. 2) Insert a USB flash memory with 10 kBytes free memory left. 3) Remove write protection or a USB flash memory with 10 kBytes free memory left.
4316 | APOWER_HW_VERSION | APower board hardware version error | Meaning: Wrong hardware version for APower board detected. Reason: APower board has wrong version. Cure: Check the 'Power Board' version. The compatible version is labeled with starting number A802262, e.g. 'A802262.0'. Replace 'Power Board' with one compatible to this digitizer. See spare parts list.
4317 | UNSUPPORTED_PRODUCT_SUBTYPE | The product subtype is unsupported. | Meaning: The product subtype read from the device eLabel is not supported by this software. Reason: 1. This is not a CR 10-X (subtype 100, 110) or CR 12-X (subtype 200, 210). 2. A wrong product subtype is stored in the device eLabel. Cure: Check if this digitizer is really a CR 10-X or CR 12-X. The name plate must show the subtype 5151/1xx or 5151/2xx. This software runs only on CR 10-X or CR 12-X, but not on other digitizer types like e.g. CR 15-X (subtype 3xx).
4318 | READ_CASS_BY_NX_FAILED_NO_IP | Reading the cassette by NX failed because the IP was not found. | Meaning: Reading the cassette data by the NX failed because the IP could not be found. Reason: 1. Cassette is empty, i.e. it does not contain an IP. 2. Mechanical problem, IP cannot be transported from the cassette to the scan drum. Cure: 1. Check if the cassette contains an IP. 2. Check the infocounter. Additional information about the root cause is stored in Ithe infocounter error history.
4319 | TRANSPORT_IP_TO_SCANNER_FAILURE_NO_IP | Transport of the IP to the scanner was not successful. | Meaning: The IP was not transported to the scanner. The automatic clean-up cycle did not find an IP in the digitizer. Reason: 1. Cassette is empty, i.e. it does not contain an IP. 2. Mechanical problem, IP cannot be transported from the cassette to the scan drum. Cure: 1. Check if the cassette contains an IP. 2. Check that the cassette doesn't clamp the IP. i.e. that the IP can slide easily out of the cassette. 3. Use another cassette. 4. Check the IP Detection Light Sensor. 5. Check the transport mechanics (IP-Drawer module, Slow Scan module).
4320 | TRANSPORT_IP_TO_BEGIN_OF_ERASURE_FAILURE_NO_IP | Transport of the IP to begin of erasure was not successful. | Meaning: The IP was not transported to the begin of erasure. The automatic clean-up cycle did not find an IP in the digitizer. Reason: 1. Cassette is empty, i.e. it does not contain an IP. 2. Mechanical problem, IP cannot be transported from the cassette to the scan drum. Cure: 1. Check if the cassette contains an IP. 2. Check that the cassette doesn't clamp the IP. i.e. that the IP can slide easily out of the cassette. 3. Use another cassette. 4. Check the IP Detection Light Sensor. 5. Check the transport mechanics (IP-Drawer module, Slow Scan module).
4321 | PREVENTIVE_MAINTENANCE_CONFIRM_EXPIRED | The preventive maintenance interval has expired. | Meaning: The preventive maintenance interval has expired. Reason: The configured interval for preventive maintenance has expired. Cure: Perform the preventive maintenance.
4322 | PREVENTIVE_MAINTENANCE_EXPIRED | The preventive maintenance interval has expired. | Meaning: The preventive maintenance interval has expired. Reason: The configured interval for preventive maintenance has expired. Cure: Perform the preventive maintenance.
4323 | MAINTENANCE_CONFIRM_REPLACE_CASSETTE_UNIT | Replace the cassette unit due to end of its life time. | Meaning: The maintenance plan requests to replace the cassette unit. Reason: The maximum life time of the cassette unit is reached. Cure: Replace the cassette unit.
4324 | MAINTENANCE_REPLACE_CASSETTE_UNIT | Replace the cassette unit due to end of its life time. | Meaning: The maintenance plan requests to replace the cassette unit. Reason: The maximum life time of the cassette unit is reached. Cure: Replace the cassette unit.
4325 | MAINTENANCE_CONFIRM_REPLACE_SLOWSCAN_DRUM | Replace the slow scan drum with pulley due to end of its life time. | Meaning: The maintenance plan requests to replace the slow scan drum with pulley. Reason: The maximum life time of the slow scan drum with pulley is reached. Cure: Replace the slow scan drum with pulley.
4326 | MAINTENANCE_REPLACE_SLOWSCAN_DRUM | Replace the slow scan drum with pulley due to end of its life time. | Meaning: The maintenance plan requests to replace the slow scan drum with pulley. Reason: The maximum life time of the slow scan drum with pulley is reached. Cure: Replace the slow scan drum with pulley.
4327 | MAINTENANCE_CONFIRM_REPLACE_SLOWSCAN_PLATE | Replace the slow scan plate with belt tensioner due to end of its life time. | Meaning: The maintenance plan requests to replace the slow scan plate with belt tensioner. Reason: The maximum life time of the slow scan plate with belt tensioner is reached. Cure: Replace the slow scan plate with belt tensioner.
4328 | MAINTENANCE_REPLACE_SLOWSCAN_PLATE | Replace the slow scan plate with belt tensioner due to end of its life time. | Meaning: The maintenance plan requests to replace the slow scan plate with belt tensioner. Reason: The maximum life time of the slow scan plate with belt tensioner is reached. Cure: Replace the slow scan plate with belt tensioner.
4329 | MAINTENANCE_CONFIRM_REPLACE_IP_DRAWER_MODULE | Replace the IP drawer module due to end of its life time. | Meaning: The maintenance plan requests to replace the IP drawer module. Reason: The maximum life time of the IP drawer module is reached. Cure: Replace the IP drawer module.
4330 | MAINTENANCE_REPLACE_IP_DRAWER_MODULE | Replace the IP drawer module due to end of its life time. | Meaning: The maintenance plan requests to replace the IP drawer module. Reason: The maximum life time of the IP drawer module is reached. Cure: Replace the IP drawer module.
4331 | MAINTENANCE_CONFIRM_LUBRICATE_CASSETTE_UNIT | Lubricate the cassette unit. | Meaning: The maintenance plan requests the lubrication of the cassette unit. Reason: The maximum defined cycles for lubrication of the cassette unit are reached. Cure: Lubricate the cassette unit according to the maintenance instruction.
4332 | MAINTENANCE_LUBRICATE_CASSETTE_UNIT | Lubricate the cassette unit. | Meaning: The maintenance plan requests the lubrication of the cassette unit. Reason: The maximum defined cycles for lubrication of the cassette unit are reached. Cure: Lubricate the cassette unit according to the maintenance instruction.
4333 | SERIAL_ID_FILE_CORRUPTED | Problem writing serial.id | There is s problem with writing the file c:/serial.id. The file could not be written or is corrupted
4334 | COPY_IDDATA_FAILED | Synchronize data failed | Product ID data d:/PMIProductData.xml is corrputed. Restore of product ID Data from USB failed, copying file failed
4335 | COPY_CONFIG_FAILED | Synchronize data failed | Copying configuration data from USB to PMI flash failed.
4336 | COPY_NOENVFILE_SPECIFIED | Synchronize data failed | No filename specified when trying to setup environment
4337 | COPY_CANNOT_FIND_NETUP | Synchronize data failed | No NET.UP found when trying to set environment
4338 | MAINTENANCE_CONFIRM_REPLACE_COVER_CASSETTE_UNIT | Replace the cover cassette unit due to end of its life time. | Meaning: The maintenance plan requests to replace the cover cassette unit. Reason: The maximum life time of the cover cassette unit is reached. Cure: Replace the cover cassette unit.
4339 | MAINTENANCE_REPLACE_COVER_CASSETTE_UNIT | Replace the cover cassette unit due to end of its life time. | Meaning: The maintenance plan requests to replace the cover cassette unit. Reason: The maximum life time of the cover cassette unit is reached. Cure: Replace the cover cassette unit.
7168 | CAS_LATENCY_NOT_SUPPORTED | FPGA initialization failed because CAS_LATENCY is not supported | Meaning: FpgaAccess: FPGA initialization failed because CAS_LATENCY is not supported. Reason: SW error. Cure: Install new software; inform R&amp;D team.
7169 | BURST_LENGTH_NOT_SUPPORTED | FPGA initialization failed because BURST_LENGTH is not supported | Meaning: FpgaAccess: FPGA initialization failed because BURST_LENGTH is not supported. Reason: SW error. Cure: Install new software; inform R&amp;D team.
7170 | SDRAM_CONTROLLER_START_FAILED | FPGA initialization failed because SDRAM controller do not start | Meaning: FpgaAccess: FPGA initialization failed because SDRAM controller do not start. Reason: HW error. Cure: Replace PMI board.
7171 | TEST_FPGA_NOT_LOADED | FPGA test failed FPGA is not loaded yet | Meaning: FpgaAccess: FPGA test failed FPGA is not loaded yet. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; inform R&amp;D team in any case.
7172 | PATTGEN_FPGA_NOT_LOADED | FPGA pattern generation failed because FPGA is not loaded yet | Meaning: FpgaAccess: FPGA pattern generation failed because FPGA is not loaded yet. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; inform R&amp;D team in any case.
7173 | PATTGEN_FLUSH_FIFO_UNDERRUN | FPGA pattern generation failed because flush FIFO underrun occured | Meaning: FpgaAccess: FPGA pattern generation failed because flush FIFO underrun occured. Reason: HW error. Cure: Replace PMI board.
7174 | PATTGEN_FLUSH_FIFO_OVERFLOW | FPGA pattern generation failed because flush FIFO overflow occured | Meaning: FpgaAccess: FPGA pattern generation failed because flush FIFO overflow occured. Reason: HW error. Cure: Replace PMI board.
7175 | PATTGEN_FLUSH_FIFO_DRAIN_NOK | FPGA pattern generation failed because flush FIFO cannot be drained | Meaning: FpgaAccess: FPGA pattern generation failed because flush FIFO cannot be drained. Reason: HW error. Cure: Replace PMI board.
7176 | FLUSH_FIFO_DRAIN_FAILED | flush FIFO cannot be drained; HW problem, fatal error | Meaning: FpgaAccess: flush FIFO cannot be drained; HW problem, fatal error. Reason: HW error. Cure: Replace PMI board.
7177 | FPGA_TEST_FAILED | FPGA test failed (one of the RAMs was not correctly accessible), fatal error | Meaning: FpgaAccess: FPGA test failed (one of the RAMs was not correctly accessible). Reason: HW error. Cure: Replace PMI board.
7178 | TIMEOUT_FILLING_SCAN_FIFO | scan fifo write pump did not fill scan fifo with SDRAM data, fatal error | Meaning: FpgaAccess: scan fifo write pump did not fill scan fifo with SDRAM data. Reason: HW error. Cure: Replace PMI board.
7179 | UNKNOWN_FLUSH_FIFO_INPUT | unknown input for flush fifo selected, fatal software error | Meaning: FpgaAccess: unknown input for flush fifo selected. Reason: SW error. Cure: Install new software; inform R&amp;D team.
7180 | FILLSDRAM_FPGA_NOT_LOADED | filling the SDRAM failed because FPGA is not loaded yet | Meaning: FpgaAccess: filling the SDRAM failed because FPGA is not loaded yet. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; inform R&amp;D team in any case.
7181 | FILLSDRAM_IMAGE_TOO_BIG | filling the SDRAM failed because desired image is too big | Meaning: FpgaAccess: filling the SDRAM failed because desired image is too big. Reason: SW error. Cure: Install new software; inform R&amp;D team.
7182 | PATTGEN_SDRAM_SIZE_UNKNOWN | FPGA pattern generation failed because size of SDRAM is not determined yet (call initFpga first!) | Meaning: FpgaAccess: FPGA pattern generation failed because size of SDRAM is not determined yet. Reason: SW error (call initFpga first!). Cure: Install new software; inform R&amp;D team.
7183 | PATTGEN_EXCEEDS_SDRAM | FPGA pattern generation failed because size of requested pattern doesn't fit into SDRAM | Meaning: FpgaAccess: FPGA pattern generation failed because size of requested pattern doesn't fit into SDRAM. Reason: SW error. Cure: Install new software; inform R&amp;D team.
7184 | LOAD_FPGA_CONFIGURATION | Can't load FPGA configuration, fatal error | Meaning: FpgaAccess: Can't load FPGA configuration. Reason: Hardware defect or software mismatch. Cure: Upload PMI software and FPGA configuration; if still NOK, replace PMI board
7185 | LOAD_FPGA_HARDWARE | Load FPGA: can't set FPGA into configuration mode, fatal error | Meaning: FpgaAccess: Load FPGA can't set FPGA into configuration mode. Reason: HW error. Cure: Replace PMI board.
7186 | STORED_FPGA_CONFIGURATION | No valid FPGA version found on PMI | Meaning: FPGA version doesn't fit to hard- or software version. Reason: SW or HW error. Cure: Update software including FPGA configuration; if still NOK, replace PMI board
7187 | READ_PMI_CPU | Can't read from FPGA CPU | Meaning: Can't read from FPGA CPU. Reason: Hardware error on PMI. Cure: Retry switch digitizer off/on; if still NOK, replace PMI board.
7188 | CAL_ERROR_FLAG | PMI CPU Error flag is set after geometric calibration | Meaning: Geometric calibration failed. Reason: Polygon or Laser or PMI board defective. Cure: Retry switch digitizer off/on, replace lens assembly CR30X, replace PMI board.
7189 | CAL_EOL | EOL error counter is set after geometric calibration | Meaning: Geometric calibration failed. Reason: 1. Optical path is partly interrupted.. 2. For CR30-X 1st/2nd, CR30-Xm and CRxFlex1/2 only: Calibration board is defective. Cure: 1. Clean scan line with antistatic brush, check mounting of light guide, switch digitizer off/on. 2. For CR30-X 1st/2nd, CR30-Xm and CRxFlex1/2 only: Replace calibration plate.
7190 | CAL_CAL_SENSORS | Calibration sensor error counter is set after geometric calibration | Error only applicable for CR30-X 1st/2nd, CR30-Xm and CRxFlex1/2: Meaning: Geometric calibration failed. Reason: Optical path is partly interrupted, calibration board is defective. Cure: Clean scan line with antistatic brush, switch digitizer off/on, replace calibration plate
7680 | IMAGE_SERVER_CREATION_FAILED | ImageTransferControl cannot create image data server; fatal! | Meaning: ImageTransferControl cannot create image data server. Reason: SW error. Cure: Install new software.
7681 | INVALID_MAX_PACKET_SIZE | ImageTransferControl received wrong FW packet size from PC; fatal! | Meaning: ImageTransferControl received wrong FW packet size from PC. Reason: SW error. Cure: Install new software; inform R&amp;D team.
7682 | INVALID_FW_SPEED_CODE | ImageTransferControl determined invalid FW speed; fatal! | Meaning: ImageTransferControl determined invalid FW speed. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; if still NOK, replace FW-board on PC; inform R&amp;D team in any case.
7683 | BLOCK_TRANSFER_TIMEOUT_EXCEEDED | ImageTransferControl: block transfer takes too long; either FW cable was disconnected during the block transfer or a fatal HW error occured! | Meaning: ImageTransferControl: block transfer takes too long. Reason: The transmission of an image data block takes too long. Cure: Check if FW cable is connected correctly; if still NOK, replace FW cable; if still NOK, replace PMI board; if still NOK, replace FW-board on PC.
7684 | INVALID_PC_NODE_ID | ImageTransferControl: block transfer failed because PC's node ID is invalid; fatal error! | Meaning: ImageTransferControl: block transfer failed because PC's node ID is invalid. Reason: SW error. Cure: Install new software; inform R&amp;D team.
7685 | SCAN_FIFO_OVERFLOW | ImageTransferControl: scan fifo overflow during block transfer; fatal! | Meaning: ImageTransferControl: scan fifo overflow during block transfer. Reason: HW error. Cure: Replace PMI board.
7686 | SCAN_FIFO_UNDERRUN | ImageTransferControl: scan fifo underrun during block transfer; fatal! | Meaning: ImageTransferControl: scan fifo underrun during block transfer. Reason: HW error. Cure: Problem with the scanning unit (e.g. BeginOfLine/EndOfLine detection problem); check if acrylic light guide and PMT are mounted properly; check supply voltages; check connection between PMI board and PMT/lens assembly; if OK, replace PMI board; if still NOK, replace PMT and light guide; if still NOK, replace lens assembly
7687 | FPGA_BUS_ERROR | ImageTransferControl: FPGA bus error occured; fatal! | Meaning: ImageTransferControl: FPGA bus error occured. Reason: HW error. Cure: Replace PMI board.
7688 | INVALID_NR_OF_REQUESTED_BYTES | ImageTransferControl: invalid number of requested bytes ( <=0); fatal SW bug! | Meaning: ImageTransferControl: invalid number of requested bytes ( &lt;=0). Reason: SW error. Cure: Install new software; inform R&amp;D team.
7689 | TIMEOUT_GET_PACKET_SIZE | ImageTransferControl: timeout: cannot determine used FW packet size; fatal! | Meaning: ImageTransferControl determined invalid FW speed. Reason: SW or HW error. Cure: Install new software; if still NOK, replace PMI board; if still NOK, replace FW-board on PC; inform R&amp;D team in any case.
7690 | TIMEOUT_DATA_NOT_AVAILABLE | ImageTransferControl: timeout: scanner didn't store enough data in SDRAM yet to start block transfer; wait, then try again! | Meaning: ImageTransferControl: timeout. Reason: Scanner didn't store enough data in SDRAM yet to start block transfer. Cure: No problem, PC software has to wait and try again.
7691 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump. Reason: HW error (DM (DataMover) pump or LLC (LinkLayerController) failed). Cure: Replace PMI board.
7692 | RECOVER_PACKET_FAILED | ImageTransferControl: recovering the FW packet failed; fatal error! | Meaning: ImageTransferControl: recovering the FW packet failed. Reason: HW error. Cure: Replace PMI board.
7693 | TOO_MANY_PACKETS_PER_BLOCK | ImageTransferControl: requested block is too big (exceeds max. possible number of packets per block); request smaller block! | Meaning: ImageTransferControl: requested block is too big (exceeds max. possible number of packets per block). Reason: SW error (PC must request smaller block!). Cure: Install new software; inform R&amp;D team.
7694 | FLUSH_FIFO_DRAIN_FAILED | flush FIFO cannot be drained; HW problem, fatal error! | Meaning: ImageTransferControl: flush FIFO cannot be drained. Reason: HW error. Cure: Replace PMI board.
7695 | PIXELVIEWER_FLUSH_FIFO_OVERFLOW | flush FIFO overflow occured in PixelViewer mode; HW problem, fatal error! | Meaning: ImageTransferControl: flush FIFO overflow occured in PixelViewer mode. Reason: HW error. Cure: Replace PMI board.
7696 | PIXELVIEWER_FLUSH_FIFO_UNDERRUN | flush FIFO underrun occured in PixelViewer mode; HW problem, fatal error! | Meaning: ImageTransferControl: flush FIFO underrun occured in PixelViewer mode. Reason: HW error. Cure: Replace PMI board.
7697 | PIXELVIEWER_CRC1_AND_2_NOT_EQUAL | after scanning in PixelViewer mode, the CRC1 and 2 are not equal; HW problem, fatal error! | Meaning: ImageTransferControl: after scanning in PixelViewer mode, the CRC1 and 2 are not equal. Reason: HW error. Cure: Replace PMI board.
7698 | PIXELVIEWER_START_NEXT_SCAN_FAILED | starting next scan in PixelViewer mode failed, because scanning is not enabled or is still running; HW failure (e.g. BOL missing) or SW synchronization problem, fatal error! | Meaning: ImageTransferControl: starting next scan in PixelViewer mode failed, because scanning is not enabled or is still running. Reason: HW failure (e.g. BOL missing) or SW synchronization problem. Cure: Install new software; if still NOK, replace PMI board; inform R&amp;D team in any case.
7699 | DM_ERROR_HANDLER_SETUP_FAILED | registering the callback function for DM errors failed | Meaning: ImageTransferControl: registering the callback function for DM errors failed. Reason: SW error. Cure: Install new software; inform R&amp;D team.
7700 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED_START | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed at start of block transfer; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed at start of block transfer. Reason: HW error. Cure: Replace PMI board.
7701 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED_OVFL | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed after scan fifo overflow; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed after scan fifo overflow. Reason: HW error. Cure: Replace PMI board.
7702 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED_URUN | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed after scan fifo underrun; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed after scan fifo underrun. Reason: HW error. Cure: Replace PMI board.
7703 | ENABLE_SCAN_FIFO_PUMPS_LLC_FAILED_RETRANS | ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed at start of packet retransmission; fatal error! | Meaning: ImageTransferControl: enabling the scan fifo write pump, DM pump or LLC failed at start of packet retransmission. Reason: HW error. Cure: Replace PMI board.
7704 | FW_DISCONNECTED_DURING_BLOCK_TRANSFER | ImageTransferControl: block transfer stopped because FW cable was disconnected! Image block can be retransmitted after reconnecting the FW cable. | Meaning: ImageTransferControl: block transfer stopped because FW cable was disconnected! Reason: FW cable was disconnected during transmission of image data to PC. Cure: Connect FW cable again; PC will request the image trasnmission automatically again, image is not lost.
7705 | FW_DISCONNECTED_DURING_BLOCK_TRANSFER_START | ImageTransferControl: block transfer stopped because FW cable was disconnected! Image block can be retransmitted after reconnecting the FW cable. Actual transfer was not started yet. | Meaning: ImageTransferControl: block transfer stopped because FW cable was disconnected! Reason: FW cable was disconnected during transmission of image data to PC. Cure: Connect FW cable again; PC will request the image trasnmission automatically again, image is not lost.
7706 | FW_DISCONNECTED_DURING_BLOCK_TRANSFER_PUMP | ImageTransferControl: block transfer stopped because FW cable was disconnected! Image block can be retransmitted after reconnecting the FW cable. DM pump was not enabled yet. | Meaning: ImageTransferControl: block transfer stopped because FW cable was disconnected! Reason: FW cable was disconnected during transmission of image data to PC. Cure: Connect FW cable again; PC will request the image transmission automatically again, image is not lost.
7707 | IMAGE_BUFFER_CREATION_FAILED | ImageTransferUDPControl: the DRAM buffer to hold the image coming from the FPGA, cannot be created. | Meaning: ImageTransferUDPControl: the DRAM buffer to hold the image coming from the FPGA, cannot be created. Reason: not enough memory resources (probably software bug) Cure: Install new software.
7708 | MORE_DATA_REQUESTED_THAN_SCANNED | ImageTransferUDPControl: more data requested than scanned | Meaning: ImageTransferUDPControl: more data requested than scanned. Reason: Index out of range. Cure: Install new software; inform R&amp;D team.
7709 | SERVER_DISCONNECTED_DURING_BLOCK_TRANSFER | ImageTransferUDPControl: block transfer stopped because server was disconnected! | Meaning: ImageTransferUDPControl: block transfer stopped because server was disconnected! Reason: Cable was disconnected during transmission of image data to PC. Cure: Connect cable again; PC will request the image trasnmission automatically again, image is not lost.
7710 | NO_DATA_IN_DRAM_TIMEOUT | ImageTransferUDPControl: block transfer stopped because no data was copied on time from FPGA into DRAM. | Meaning: ImageTransferUDPControl: block transfer stopped because no data was copied on time from FPGA into DRAM. Reason: Copying data into DRAM takes too long, or, no data anymore in FPGA. Cure: Install new software; inform R&amp;D team.
7711 | DMA_NOT_PREPARED_TIMEOUT | ImageTransferUDPControl: Difo already wants to start transferring the first block while the DMA process is not even ready to start filling RAM. | Meaning: ImageTransferUDPControl: Difo already wants to start transferring the first block while the DMA process is not even ready to start filling RAM. Reason: Timing problem. Cure: Install new software; inform R&amp;D team.
7712 | PATGEN_ENABLE_PROBLEM | ImageTransferUDPControl: The Pattern Generator cannot be enabled | Meaning: ImageTransferUDPControl: The Pattern Generator cannot be enabled. Reason: Could be software or hardware problem. Cure: Install new software; inform R&amp;D team.
7713 | PATGEN_DISABLE_PROBLEM | ImageTransferUDPControl: The Pattern Generator cannot be disabled | Meaning: ImageTransferUDPControl: The Pattern Generator cannot be disabled. Reason: Could be software or hardware problem. Cure: Install new software; inform R&amp;D team.
7714 | IMAGE_TRANSFER_CANCELLED | ImageTransferUDPControl: the image transfer was cancelled | Meaning: ImageTransferUDPControl: the image transfer was cancelled.
7715 | DM_ERROR | ImageTransferControl: block transfer stopped because the DataMover (DM) hardware reported an error | Meaning: ImageTransferControl: block transfer stopped because the DataMover (DM) hardware reported an error. Reason: A DM error occurred during the transmission of an image block. Cure: The PC will automatically request the image block again (up to 10 times); if all repetitions fail an image transer error is displayed; in the latter case reboot the PC, the PC will start the image transmission automatically again, the image is not lost as long as the digitizer is not switched off.
10752 | INCIDENCE_OF_LIGHT | Incidence of light --> adj.PMM/PMT failed ; fatal! | Meaning: It is too bright to adjust the Photo Multiplier. Reason: Housing of digitizer is open. Cure: Close the digitizer housing.
10753 | PMM_ADJUST | PMM/PMT adjust failed | Meaning: Photo multiplier malfunction. Reason: - Cure: Switch digitizer off/on and try a scan, if failed again replace PMT + LIGHT COLLECTOR.
10754 | GEO_CALIBRATION | Geometric and jitter calibration failed | Meaning: Geometric and jitter calibration failed. Reason: Optic module, sensor path misadjusted. Cure: Check proper seating of optic module, clean scan path with dust brush.
10755 | START_POLYGON | Start polygon failed | --
10756 | ADC_ADJUST | ADC adjustment out of spec | Meaning: ADC offset adjust out of range. Reason: ADC defective. Cure: Switch digitizer off/on, replace PMI board.
10757 | START_SCAN | Start scan failed | Meaning: Start scan failed, because scan is already running;Reason: Software error;Cure: Update software or inform RDE
10758 | POLYGON_LOCKED | Polygon get not locked | Meaning: Polygon can't reach the requested speed. Reason: Polygon defective. Cure: Replace optic unit.
10759 | PREPARE_SCAN | Prepare Scan failed | --
10760 | TOO_LESS_DATA | Too less data stored in SDRAM --> SDRAM filled with overexposed pixel values | Meaning: Too less data stored in SDRAM --&gt; SDRAM filled with overexposed pixel values. Reason: Electrostatic discharge, error in begin or end of line signal (BOL,EOL). Cure: 1. Restart the Digitizer and check, if the problem persists. 2. Replace PMI board. 2. Replace Optic Module.
10761 | TEST_SCAN | Test scan failed --> BOL or EOL out of range | Obsolete
10762 | PMM_REFCAL | Calculation of REFCAL from IPF out of range | Meaning: The IP depended calculation of sensitivity factor results in a value out of range. Reason: Wrong PMM/PMT adjustment values stored in PMT eLabel. Cure: Replace PMT + LIGHT COLLECTOR.
10763 | PMM_AGING_OBSOLETE | PMT sensitivity become critical --> reaches 94= 750V | Meaning: The PMTs end is coming soon. Cure: Replace PMT + LIGHT COLLECTOR
10764 | PMM_FINE_ADJUST_OBSOLETE | PMM/PMT fine adjust failed --> possibly bad image quality | Meaning: PMT sensitivity adjustment failed Reason: - PMI POWER DISTRIBUTION board defective - PMT defective or end of life Cure: 1. Replace PMI POWER DISTRIBUTION board 2. If error still persists replace PMT + LIGHT COLLECTOR
10765 | START_GEO_CALIBRATION | Start geometric and jitter calibration failed | --
10766 | IPF_OUT_OF_RANGE_OBSOLETE | IP sensitivity factor out of range | Meaning: IP is not compatible with digitizer Reason: IP does nit fit to this digitizer Cure: Use an IP, which is compatible to digitizer
10767 | TEST_SCAN_CAL_OBSOLETE | Test scan failed: Not all CAL sensors detected --> defect or IP in digitizer | Meaning: Can't detect the expected amount of calibration signals. Reason: Optic path interrupted, IP in digitizer, broken light collector, cable PMI to calibration board defective, CAL-board defective Cure: Clean the scan line, check optic path, check cable PMI to calibration board, replace calibration board.
10768 | READ_PMM_ELABEL_OBSOLETE | Can't read PMM settings from PMM eLabel | --
10769 | READY_TO_SCAN_CAL_SENSORS_OBSOLETE | Calibration sensor defective or two IP's in scanner | Meaning: Can't detect the expected amount of calibration signals. Reason: CAL-board defective or optic path interrupted (IP or dust). Cure: Clean the scan line, replace calibration board.
10770 | PMM_BLUE_LED | Photo Multiplier Module (PMM/PMT), blue LED or Power Board defective | Meaning: Can't detect PMT blue LED. Reason: - PMT blue LED defective or not connected - PMI Power Board defective - PMT defective Cure: 1. Check connection and cable from blue LED to PMT 2. Replace PMI Power Board 3. If error still persists replace PMT + LIGHT COLLECTOR
10771 | FPGA_BUS_OBSOLETE | FPGA BUS error detected | Meaning: FPGA BUS error occurred. Reason: unknown Cure: Restart digitizer, if error occurred again replace PMI board.
10772 | AUTOEOL_ADDED_OBSOLETE | FPGA has added an EOL at end of EOL valid | Meaning: No EOL detected, FPGA adds one at end of EOL valid window. Reason: Galvo or optics error. Cure: Restart digitizer, replace Optic Module.
10773 | PMM_OFFSET_VOLTAGE_OBSOLETE | PMM/PMT offset voltage out of range | Meaning: PMT offset voltage adjustment failed Reason: - Incidence of light - PMI POWER DISTRIBUTION board defective - PMT defective (end of life) Cure: 1. Make sure all covers are closed 2. Replace PMI POWER DISTRIBUTION board 3. If error still persists replace PMT + LIGHT COLLECTOR
10774 | IMAGE_DMA | Image transfer error | Meaning: Image transfer failed Reason: - Hard- or Software error Cure: 1. Save Logfile and send it to RDE 2. Restart the Digitizer and check, if the problem persists
10775 | TOO_LESS_DATA_WARNING_OBSOLETE | Too less data stored in SDRAM --> SDRAM filled with RAMP | Meaning: Too less data stored in SDRAM --&gt; SDRAM filled with stripes. Reason: Electrostatic discharge, error in begin or end of line signal (BOL,EOL). Cure: Check grounding of calibration board and grounding of front cover.
10776 | PMM_SPI_ACCESS | Access to PMM SPI-BUS failed | Meaning: Can not get PMM SPI-BUS in time. Reason: -. Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check the connector PMI - PMT 3. If error still persists replace PMT + LIGHT COLLECTOR
10777 | PMM_REFLUM | Adjust RefLum failed | Meaning: Can not adjust blue LED for the needed brightness. Reason: - blue LED defective - Phototransistor defective Cure: 1. Restart the Digitizer and check, if the problem persists 2. If error still persists replace PMT + LIGHT COLLECTOR
10778 | IMAGE_DMA_OPEN | Image transfer error | Error during image scan. Meaning: The scan process didn't start successful. Initializing image transfer inside Digitizer memory failed. Reason 1: Software error. Cure 1: Restart digitizer to return IP into cassette and release cassette. At the restart you get relevant selftest result. Install the newest digitizer software. Reason 2: Hardware error. 'PMI Board' defective. Cure 2. Replace 'PMI Board'.
10779 | IMAGE_DMA_PEND | Image transfer error | Error during image scan. Meaning: The scan process didn't start successful. Function sc_pend at beginning of image transfer inside Digitizer memory failed. Reason 1: Software error. Cure 1: Restart digitizer to return IP into cassette and release cassette. At the restart you get relevant selftest result. Install the newest digitizer software. Reason 2: Hardware error. 'PMI Board' defective. Cure 2. Replace 'PMI Board'.
10780 | UNKNOWN_APPTYPE | Got unknown application type | Prepare Scan error. Meaning: Got unknown application type from the image processing software. The received data for application type is corrupt or not supported at this device. Reason 1: Configuration error. Cure 1: Configure the workstation for the right device. Reason 2: Software error at the image processing software. Cure 2: Reinstall the image processing software. Configure the workstation for the right device.
10852 | GALVO_INIT | Adjust Galvo frequency failed | Meaning: Can not adjust Galvo frequency. Reason: - Galvo defective - Laser beam disadjusted Cure: 1. Restart the Digitizer and check, if the problem persists 2. If error still persists replace optic module
10853 | BOS_NO_DATA | Begin of scan no data available | Meaning: Galvo or Laser defective Reason: - Galvo defective - Laser defective Cure: 1. Restart the Digitizer and check, if the problem persists 2. If error still persists replace optic module
10854 | DECODE_BARCODE | Decode barcode failed | Meaning: Can't detect barcode Reason: - Barcode or IP defective - No IP in scanner - Optic module defective - PMM defective Cure: 1. Restart the Digitizer and check, if the problem persists 2. Try it again with another cassette 3. If error still persists replace PMM 4. Replace optic module
10855 | DETECT_BARCODE | Detecting of barcode in on IP failed or no IP | Meaning: Detecting barcode label failed. Reason: IP not in scanner or no barcode label on IP. Cure: 1. Try it again with another cassette 2. Check an clean barcode on IP 3. ... 4. If error still persists replace optic module
10856 | BARCODE_NO_DATA | Got no data for barcode read | Meaning: Galvo or Laser defective Reason: - Galvo defective - Laser defective Cure: 1. Restart the Digitizer and check, if the problem persists 2. If error still persists replace optic module
10857 | SCAN_BOL_EOL_MISSING | During scan a BOL or EOL signal is missing | Meaning: A BOL (BeginOfLine) or EOL (EndOfLine) signal does not arrive in time. Reason: 1. The Optic Module synchronisation failed because of mechanical shock. 2. The newest digitizer software is not installed. 3. The cable to Optic Module is defective or disconnected. 4. The Optic Module is defective. Cure: 1. Restart the Digitizer and check, if the problem persists. 2. Install the newest digitizer software. 3. Check cable PMI Board to Optic Module. 4. Replace Optic Module.
10858 | GALVO_ADJUST_DRIVE_PULSE_TOO_HIGH | Adjust Galvo line length failed | Meaning: Can not adjust Galvo line length. Reason: - Galvo defective Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check if ambient air temperature and relative humidity in spec. 3. If error still persists replace optic module
10859 | GALVO_ADJUST_LINE_LENGTH | Adjust Galvo line length failed | Meaning: Can not adjust Galvo line length. Reason: - Galvo defective Cure: 1. Restart the Digitizer and check, if the problem persists 2. If error still persists replace optic module
10860 | MISSING_LINES_IN_IMAGE | Single lines are missing in the scanned image. | Meaning: Single lines are missing in the scanned image without impact on image quality. Reason: During scan the signal BOL (BeginOfLine) or EOL (EndOfLine) was not detected for one line. Therefore single scanned lines were not stored as image data. Cure: No action required. If the problem appears frequently, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
10861 | TOO_MANY_MISSING_LINES_IN_IMAGE | Too many lines are missing in the scanned image. | Meaning: Too many lines are missing in the scanned image, which may have an impact on image quality. Reason: During scan the signal BOL (BeginOfLine) or EOL (EndOfLine) was not detected for several lines. Therefore several scanned lines were not stored as image data. Cure: 1. Check if the installation site requirements (shock and vibration, power failure, potential difference and electromagnetic interference) are fulfilled. 2. Check all cables between optics module and PMI board. 3. If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
11264 | PC_TO_DIC_ADAPTER_RECEIVE_DEFAULT_NOT_POSSIBLE | Adapter tried to receive a message using the 'receiveDefault' method, but it contains unsupported number of bytes | --
11265 | BASE_APP_SEND_FAILED_NO_CONNECTION | RTPort_send failed because connection to receiver is not established | --
11266 | BASE_APP_SEND_FAILED_NO_MEMORY | RTPort_send failed because no memory available (list of unused messageobjects exhausted) | --
11267 | BASE_APP_SEND_FAILED_INVALID_PRIO | RTPort_send failed because message priority is invalid | --
11268 | BASE_APP_SEND_FAILED_OTHER_REASON | RTPort_send failed for any other reason | --
14848 | READ_OPTICS_TYPE | Can't read optics type from optics eLabel -> using default type | Meaning: Default value for optic type are used. Reason: Unreadable entry in 'optic type' of optic module elabel. Cure: Replace optic module.
14849 | READ_POLYGON_TYPE | Can't read polygon type from optics eLabel -> using default type | Meaning: Default values for polygon type are used. Reason: Unreadable entry in polygon type of optic module elabel. Cure: Replace optic module.
14850 | READ_OPERATION_MODE | Can't read operation mode from PMI eLabel -> setting to normal mode | Meaning: 'Normal mode' is used as operation mode. Reason: Unreadable entry in 'operation mode' of PMI elabel. Cure: Replace PMI board.
14851 | READ_POLYGON_SPEED_OFFSET | Can't read polygon speed offset from optics eLabel -> using offset 0 | --
14852 | INIT_PMM | Initializing I2C hardware for PMM failed | Meaning: Communication error PMI to PMM. Reason: PMI and PMM connection not OK or PMI or PMM defective. Cure: 1) Check PMM cabling. 2) Replace PMM cabling. 3) Replace PMI board. 4) Replace PMM.
14853 | CAL_BOARD_CONNECTOR | Calibration board not detected | Error only applicable for CR30-X 1st/2nd, CR30-Xm and CRxFlex1/2: Meaning: Calibration board not detected. Reason: Cable broken or calibration board defective. Cure: 1) Check calibration board cabling. 2) Replace calibration board, if error persists.
14854 | OPTICS_BOARD_CONNECTOR | Optics board not detected | Meaning: Optic board not detected. Reason: Cable broken/not connected or optic board defective. Cure: 1) Check optic module cabling. 2) Replace optic module, if error persists.
14855 | LASER_DIODE | Check laser diode failed | Meaning: Check laser diode failed. Reason: Laser diode defective. Cure: 1. Restart the Digitizer and check, if the problem persists 2. Exchange the OPTIC MODULE, refer to the Service Manual for order and replacement
14856 | OPTICS_TYPE | Unknown optics type detected | Meaning: Unknown optics type detected Reason: Unknown optics type in elabel. Cure: 1. Restart the Digitizer and check, if the problem persists 2. Exchange the OPTIC MODULE, refer to the Service Manual for order and replacement
14857 | INIT_ELABEL | Error initializing optics eLabel | Meaning: Initialization of OPTIC MODULE e-label (module specific parameter storage) failed. The Digitizer is out of operation and requires service support. Reason: OPTIC MODULE e-label can't be initialized due to: - OPTIC MODULE not connected - Defective of the PMI-BOARD - OPTIC MODULE cable or connectors - Corrupted data structure on the e-label - Defective OPTIC MODULE Cure: 1. Restart the Digitizer and check, if the problem persists 2. Check cabling and connection between PMI-BOARD and OPTIC MODULE 3. Exchange the PMI-BOARD, refer to the Service Manual for order and replacement 4. Exchange the OPTIC MODULE, refer to the Service Manual for order and replacement
14858 | PMM_ELABEL | Error getting data from PMM eLabel | Meaning: Can't read data from PMM eLabel. Reason: PMM eLabel defective or PMM not connected. Cure: 1) Check PMM cabling. 2) Replace PMM cabling. 3) Replace PMI board. 4) Replace PMM.
14859 | PMM_BOARD_CONNECTOR | PMM board not detected | Meaning: PMM board not detected Reason: PMM is not connected to PMI board. Cure: 1) Check PMM cabling. 2) Replace PMM cabling, if error persists. 3) Replace PMM, if error persists.
14860 | START_POLYGON | Start or Lock polygon failed | Meaning: Polygon does not reach its speed. Reason: Power supply or polygon defective. Cure: Check power supply, replace optics module.
14861 | UPDATE_PMM | Error: PMM firmware update failed - Switch digitizer Off/On for retry | Meaning: PMM software update failed. Reason: PMM software cannot be updated. Cure: 1) Check PMM cabling. 2) Replace PMM, if error persists.
14862 | UPDATE_PMM_FAILED | Error: PMM firmware update failed | Meaning: PMM software update failed. Reason: PMM software cannot be updated. Cure: 1) Check PMM cabling. 2) Replace PMM, if error persists.
14863 | PMMADJ_LPV_TOO_HIGH | Prepare PMM adjust: LPV is too high | --
14864 | PMMADJ_LED_BRIGHTNESS | Prepare PMM adjust: Error LED not bright enough | --
14865 | PMMADJ_SEL_SENS | Prepare PMM adjust: Error selecting sensitivity | --
14866 | PMMADJ_INCIDENCE_OF_LIGHT | Prepare PMM adjust: Error because incidence of light | --
14867 | PMMADJ_WRITE_ROV | Analyse PMM adjust: Error writing new REFCAL to PMM | --
14868 | PMMADJ_PARAMETER | Analyse PMM adjust: Error parameter real dose or target index out of range | --
14869 | BOL_TIMEOUT | Lost BOL signal | Meaning: ESD or optics problem. Reason: electric discharge or optics module defective. Cure: if permanent check optics module.
14870 | PMI_ELABEL | Error getting data from PMI eLabel | Meaning: Error getting data from PMI eLabel. Reason: PMI eLabel defective. Cure: Replace PMI board.
14871 | DIST_BOL_EOL_OUT_OF_RANGE | Distance BOL to EOL out of range | Meaning: Distance BOL to EOL out of range. Reason: Polygon speed out of range, Optics module out of range. Cure: Check polygon speed (image width), replace optics module.
14872 | PMI_HW_VER | Software needs PMI-Board hardware version 1 or later | Meaning: PMI software does not fit to PMI hardware. Reason: Software update to an incompatible PMI board. Cure: Replace PMI board to index 1 or higher.
14873 | PMI_FPGA_VER | FPGA version does not fit to software version | Meaning: PMI software does not fit to PMI FPGA hardware. Reason: Software update failed or not completed. Cure: Update PMI software.
14874 | OPTICS_ELABEL_CORRUPTED | Optics eLabel data corrupted | Meaning: Optics eLabel can't be read, need to be restored from PC. Reason: Cable from PMI to optics module too long or unknown. Cure: Restore optics eLabel from PC.
14875 | LD_AGING | Laser diode is aging | Meaning: Laser diode is supplied with additional electrical power (1.5 of the standard value) to compensate aging for a while. Reason: Laser diode is short before end of life.
14876 | INIDOSEDET_IPTYPEORDOSE | IP type or dose out of range | Meaning: IP type or dose out of range. Reason: Software error. Cure: Install correct software.
14877 | INIDOSEDET_WRONGHARDWARE | initOverExpComparatorsFromData called on PMI1/2 | Meaning: The function initOverExpComparatorsFromData was called on the wrong hardware PMI1 or 2. Reason: Software error. Cure: Install correct software.
14878 | PMM_BLUE_LED | Photo Multiplier Module (PMM/PMT), blue LED defective or not connected | Meaning: Can't detect PMT blue LED. Reason: - PMT blue LED defective or not connected - PMT defective Cure: 1. Check connection and cable from blue LED to PMT 2. If error still persists replace PMT + LIGHT COLLECTOR
14879 | PMM_TUBE_BOARD | Photo Multiplier Module (PMM/PMT)defective or tube board not connected | Meaning: Can't read set HV. Reason: - PMM defective or tube board not connected Cure: 1. Check coaxial connection between PMI board and PMM tube board. 2. Check connection and cable from PMM main to PMM tube board 3. If error still persists replace PMT + LIGHT COLLECTOR
14880 | OPTICS_ELABEL | Error getting data from optics eLabel | Meaning: Can't read data from optics eLabel. Reason: Optic module eLabel defective or optic module not connected. Cure: 1) Check optic module cabling. 2) Replace optic module cabling. 3) Replace PMI board. 4) Replace optic module.
14881 | COULD_NOT_READ_PMM_FLASH_CNT | Error reading counters from PMM flash | Meaning: Cannot read counters from PMM flash memory. Cycle counter and operating hours may not correspond to real values of this PMM module.
14882 | COULD_NOT_WRITE_PMM_FLASH_CNT | Error writing counters into PMM flash | Meaning: Cannot write counters into PMM flash memory. Cycle counter and operating hours may not correspond to real values of this PMM module.
14883 | COULD_NOT_WRITE_OPTICS_ELABEL_CNT | Error writing counters into optics eLabel | Meaning: Cannot write counters into optics eLabel. Cycle counter and operating hours may not correspond to real values of this optics module.
14884 | COULD_NOT_READ_OPTICS_ELABEL_CNT | Error reading counters from optics eLabel | Meaning: Cannot read counters from optics eLabel. Cycle counter and operating hours may not correspond to real values of this optics module.
14885 | COULD_NOT_READ_OPTICS_ELABEL_INFO | Error reading serial number from optics eLabel | Meaning: Cannot read serial number from optics eLabel . Reason: The optics eLabel is empty, was formatted, contains invalid data or is disconnected. Cure: 1. Check cabling connection between PMI board and optics module. 2. Format the optics eLabel via ESF, then write valid data to the optics eLabel via ESF.
14886 | COULD_NOT_READ_PMM_INFO | Error reading serial number from PMM flash memory | Meaning: Cannot read serial number from PMM flash memory . Reason: The flash memory is empty, was formatted, contains invalid data or is disconnected. Cure: 1. Check cabling connection between PMI board and PMM module. 2. Write valid data to the PMM flash memory via ESF.
14887 | COULD_NOT_READ_SENS_DATA | Could not read sensitivity data | Meaning: Cannot read PMM sensisitivity data. Reason: PMM data not initialized (at production only) or PMM defective Cure: Adjust PMM (at production only) 1. Restart the Digitizer and check, if the problem persists 2. Check connection and cable from PMT to PMI 3. If error still persists replace PMT + LIGHT COLLECTOR
15360 | DIC_TO_IHC_NO_RECEIVER_FOUND | no suitable receiver found while DIC is sending to HDL | Meaning: No receiver found after sending message from DigitizerControl (DIC) on PMI-board to IP-HandlingControl (HDL) on 'Power Board'. Reason: Software error, input queue of HDL may be full. Cure: Install new software; inform R&amp;D team.
15361 | DIC_TO_IHC_UNKNOWN_CHANNEL | unknown CAN channel for sending from DIC to HDL | Meaning: The CAN channel number for sending message from DigitizerControl (DIC) on PMI-board to IP-HandlingControl (HDL) on 'Power Board' is unknown. Reason: Software error. Cure: Install new software; inform R&amp;D team.
15362 | DIC_TO_IHC_ERR_READING_ACKNOWLEDGE | DIC could not read acknowledge from HDL | Meaning: Could not read acknowledge after sending message from DigitizerControl (DIC) on PMI-board to IP-HandlingControl (HDL) on 'Power Board'. Reason: 1. CAN communication problem. 2. Software error. 3. 'Power Board' is defective. Cure: 1. Check all CAN connections and power connection on 'Power Board'. 2. Install new software. 3. Replace 'Power Board'. 4. If still not working, inform R&amp;D team.
15363 | DIC_TO_IHC_CAN_ERROR | sending from DIC to HDL failed with CAN fast channel error | Meaning: CAN error during sending message from DigitizerControl (DIC) on PMI-board to IP-HandlingControl (HDL) on 'Power Board'. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&amp;D team.
15372 | UID_TO_UIC_NO_RECEIVER_FOUND | no suitable receiver found while UID is sending to UIC | Meaning: No receiver found after sending message from UserInterfacDispatcher(UID) on PMI-board to UserInterfaceControl (UIC) on 'Power Board'. Reason: Software error, input queue of UIC may be full. Cure: Install new software; inform R&amp;D team.
15373 | UID_TO_UIC_UNKNOWN_CHANNEL | unknown CAN channel for sending from UID to UIC | Meaning: The CAN channel number for sending message from UserInterfacDispatcher(UID) on PMI-board to UserInterfaceControl (UIC) on 'Power Board' is unknown. Reason: Software error. Cure: Install new software; inform R&amp;D team.
15374 | UID_TO_UIC_ERR_READING_ACKNOWLEDGE | UID could not read acknowledge from UIC | Meaning: Could not read acknowledge after sending message from UserInterfacDispatcher(UID) on PMI-board to UserInterfaceControl (UIC) on 'Power Board'. Reason: CAN communication problem or software error. Cure: Check all CAN connections and power connection on 'Power Board' ; if connections are OK, install new software and inform R&amp;D team.
15375 | UID_TO_UIC_CAN_ERROR | sending from UID to UIC failed with CAN fast channel error | Meaning: CAN error during sending message from UserInterfacDispatcher(UID) on PMI-board to UserInterfaceControl (UIC) on 'Power Board'. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&amp;D team.
15380 | IHC_TO_DIC_NO_RECEIVER_FOUND | no suitable receiver found while HDL is sending to DIC | Meaning: No receiver found after sending message from IP-HandlingControl (HDL) on 'Power Board' to DigitizerControl (DIC) on PMI-board. Reason: Software error, input queue of DIC may be full. Cure: Install new software; inform R&amp;D team.
15381 | IHC_TO_DIC_UNKNOWN_CHANNEL | unknown CAN channel for sending from HDL to DIC | Meaning: The CAN channel number for sending message from IP-HandlingControl (HDL) on 'Power Board' to DigitizerControl (DIC) on PMI-board is unknown. Reason: Software error. Cure: Install new software; inform R&amp;D team.
15382 | IHC_TO_DIC_ERR_READING_ACKNOWLEDGE | HDL could not read acknowledge from DIC | Meaning: Could not read acknowledge after sending message from IP-HandlingControl (HDL) on 'Power Board' to DigitizerControl (DIC) on PMI-board. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&amp;D team.
15383 | IHC_TO_DIC_CAN_ERROR | sending from HDL to DIC failed with CAN fast channel error | Meaning: CAN error during sending message from IP-HandlingControl (HDL) on 'Power Board' to DigitizerControl (DIC) on PMI-board. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&amp;D team.
15392 | UIC_TO_UID_NO_RECEIVER_FOUND | no suitable receiver found while UIC is sending to UID | Meaning: No receiver found after sending message from UserInterfaceControl (UIC) on 'Power Board' to UserInterfacDispatcher(UID) on PMI-board. Reason: Software error, input queue of UID may be full. Cure: Install new software; inform R&amp;D team.
15393 | UIC_TO_UID_UNKNOWN_CHANNEL | unknown CAN channel for sending from UIC to UID | Meaning: The CAN channel number for sending message from UserInterfaceControl (UIC) on 'Power Board' to UserInterfacDispatcher(UID) on PMI-board is unknown. Reason: Software error. Cure: Install new software; inform R&amp;D team.
15394 | UIC_TO_UID_ERR_READING_ACKNOWLEDGE | UIC could not read acknowledge from UID | Meaning: Could not read acknowledge after sending message from UserInterfaceControl (UIC) on 'Power Board' to UserInterfacDispatcher(UID) on PMI-board. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&amp;D team.
15395 | UIC_TO_UID_CAN_ERROR | sending from UIC to UID failed with CAN fast channel error | Meaning: CAN error during sending message from UserInterfaceControl (UIC) on 'Power Board' to UserInterfacDispatcher(UID) on PMI-board. Reason: CAN communication problem or software error. Cure: Check CAN connections; if connections are OK, install new software and inform R&amp;D team.
19456 | TXFLDSENDMSG | transmitFolder() - sendMsg failed | --
19457 | TXFLDABORTASSOCIATION | transmitFolder() - abortAssociation() failed | --
19458 | CREATEOBJECT | assembleDICOMObject() - createObject failed | --
19459 | ADDPATNAME | assembleDICOMObject() - adding attr PatientName failed | --
19460 | ADDPATID | assembleDICOMObject() - adding attr PatientBirthDatefailed | --
19461 | ADDPATBIRTHDATE | assembleDICOMObject() - adding attr PatientName failed | --
19462 | ADDPATSEX | assembleDICOMObject() - adding attr PatientSex failed | --
19463 | ADDSTUDYINSTUID | assembleDICOMObject() - adding attr StudyInstanceUID failed | --
19464 | ADDSTUDYDATE | assembleDICOMObject() - adding attr StudyDate failed | --
19465 | ADDSTUDYTIME | assembleDICOMObject() - adding attr StudyTime failed | --
19466 | ADDACCESSIONNUMBER | assembleDICOMObject() - adding attr AccessionNumber failed | --
19467 | ADDREFPHYSICIAN | assembleDICOMObject() - adding attr ReferringPhysiciansName failed | --
19468 | ADDSTUDYID | assembleDICOMObject() - adding attr StudyID failed | --
19469 | ADDMODALITY | assembleDICOMObject() - adding attr Modality failed | --
19470 | OFFLINESERIESINSTUID | assembleDICOMObject() - adding attr PatientName failed | --
19471 | ADDSERIESNUMBER | assembleDICOMObject() - adding attr SeriesNumber failed | --
19472 | ADDLATERALITY | assembleDICOMObject() - adding attr Laterality failed | --
19473 | ADDBODYPART | assembleDICOMObject() - adding attr BodyPart failed | --
19474 | ADDPLATETYPE | assembleDICOMObject() - adding attr PlateType failed | --
19475 | ADDVIEWPOSITION | assembleDICOMObject() - adding attr ViewPosition failed | --
19476 | ADDMANUFACTURER | assembleDICOMObject() - adding attr Manufacturer failed | --
19477 | ADDMANUFACTURERMODELNAME | assembleDICOMObject() - adding attr ManufacturerModelName failed | --
19478 | ADDSERIALNUMBER | assembleDICOMObject() - adding attr SerialNumber failed | --
19479 | ADDSWVERSION | assembleDICOMObject() - adding attr SoftwareVersion failed | --
19480 | ADDIMAGETYPE | assembleDICOMObject() - adding attr ImageType failed | --
19481 | ADDACQDATE | assembleDICOMObject() - adding attr AcquisitionDate failed | --
19482 | ADDACQTIME | assembleDICOMObject() - adding attr AcquisitionTime failed | --
19483 | ADDINSTANCENUMBER | assembleDICOMObject() - adding attr InstaceNumber failed | --
19484 | ADDPATORIENT | assembleDICOMObject() - adding attr PatientOrientation failed | --
19485 | ADDIMAGECOMMENTS | assembleDICOMObject() - adding attr ImageComments failed | --
19486 | ADDPLATEID | assembleDICOMObject() - adding attr PlateID failed | --
19487 | ADDIMAGERPIXELSPACING | assembleDICOMObject() - adding attr ImagerPixelSpacing failed | --
19488 | ADDCASSETTESIZE | assembleDICOMObject() - adding attr CassetteSize failed | --
19489 | ADDEXPOSURESONPLATE | assembleDICOMObject() - adding attr ExposuresOnPlate failed | --
19490 | ADDRESCALEINTERC | assembleDICOMObject() - adding attr RescaleIntercept failed | --
19491 | ADDRESCALESLOPE | assembleDICOMObject() - adding attr RescaleSlope failed | --
19492 | ADDRESCALETYPE | assembleDICOMObject() - adding attr RescaleType failed | --
19493 | ADDPXLSAMPLES | assembleDICOMObject() - adding attr PixelsPerSample failed | --
19494 | ADDPHOTOMETRICINTERP | assembleDICOMObject() - adding attr PhotometricInterpretation failed | --
19495 | ADDROWS | assembleDICOMObject() - adding attr Rows failed | --
19496 | ADDCOLUMNS | assembleDICOMObject() - adding attr Columns failed | --
19497 | ADDALLOCEDBITS | assembleDICOMObject() - adding attr AllocatedBits failed | --
19498 | ADDSTOREDBITS | assembleDICOMObject() - adding attr StoredBits failed | --
19499 | ADDHIGHBIT | assembleDICOMObject() - adding attr HighBit failed | --
19500 | ADDPXLREPR | assembleDICOMObject() - adding attr PixelRepresentation failed | --
19501 | ADDSPECCHARSET | assembleDICOMObject() - adding attr SpecificCharacterSet failed | --
19502 | ADDINSTCREDATE | assembleDICOMObject() - adding attr InstanceCreationDate failed | --
19503 | ADDINSTCRETIME | assembleDICOMObject() - adding attr InstanceCreationTime failed | --
19504 | ADDSOPCLASSUID | assembleDICOMObject() - adding attr SOPClassUID failed | --
19505 | ADDSOPINSTUID | assembleDICOMObject() - adding attr SOPInstanceUID failed | --
19506 | ADDPRIVCREATOR | assembleDICOMObject() - adding attr PrivateCreator failed | --
19507 | ADDCASSETTEDATA | assembleDICOMObject() - adding attr CassetteData failed | --
19508 | ADDSCANDIRECTION | assembleDICOMObject() - adding attr ScanDirection failed | --
19509 | ADDIMG | addImageFile() failed | --
19510 | ADDCALLBACK | addImageCallback() failed | --
19511 | TXFLDRXMSG | failed to receive a response from the SCP | --
19512 | TXNDISRELEASE | failed to release the association | --
19513 | ADDIDSTUDYDESCRIPTION | assembleDICOMObject() - adding attr StudyDescription failed | --
19514 | ADDIDSERIESDESCRIPTION | assembleDICOMObject() - adding attr SeriesDescription failed | --
19515 | TXFINDPRESENTATIONCONTEXT | transmitFolder - unable to find the presentation context ID | --
19516 | DESTINATIONNOTFOUND | destination not configured | the given destination name has not been found in the list of known destination.
19517 | DESTINATIONNAMETOOLONG | name of the destination is too long | the name of the destination exceeds 64 characters.
19518 | DESTINATIONNOIPADDR | This is not an ip address | The ipaddress exceeds the length of 15 characters.
19519 | DESTINATIONCALLEDAETITLETOOLONG | The called AE title of the destination is too long | The AE title of the destination exceeds 64 characters.
19520 | NOUIDGENERATIONPOSSIBLE | The generation of unique UIDs cannot be guaranteed | The system time used to generate unique identifiers as required by DICOM shows is not valid. Exchange the battery of the real-time clock or specify a valid timeserver (if applicable).
19521 | DESTINATIONPRECAETITLETOOLONG | The AE title of the fast preview destination is too long | The AE title of the destination exceeds 64 characters.
19522 | TXOFFADDIMGFILE | failed to add imgfile hook | cannot add file reference to the image file into the DICOM object - image file existing?
19523 | TXONADDIMGFILE | failed to add img memory hook | cannot add hook to the image data in memory into the DICOM object - software failure!
19524 | BUILDASSOCRQ | failed to build the association request | check networking configuration (calling AE title, called AE title)
19525 | SENDASSOCRQ | failed to send association request | DTF library was not able to send the association request.
19526 | CONNBLDABORTASSOCIATION | cannot abort association | cannot abort assoc after problem when assembling the DICOM object
19527 | RECVANSWERTIMEOUT | error while waiting for the response | Error reported while waiting for the response of the DIMSE_sendMsg call. The image is queued again, rely on automatic recovery.
19528 | ADDIMAGESINACQUISITION | assembleDICOMObject() - adding attr images in acquisition failed | --
19968 | IPDATA_NOTACCESSIBLE | IPData NULL reference | The reference to IPData is NULL for SpoolerInput.
19969 | SCANPAR_NOTACCESSIBLE | ScanPar NULL reference | The reference to ScanParameters is NULL for SpoolerInput.
19970 | CREATEEMPTYQ | sc_qecreate failed for 'empty' queue | out of ressources or software bug
19971 | CREATEQUEUEDQ | sc_qecreate failed for 'queued' queue | out of ressources or software bug
19972 | CREATETXEDQ | sc_qecreate failed for 'transmitted' queue | out of ressources or software bug
19973 | CREATEWARNQ | sc_qecreate failed for 'warning' queue | out of ressources or software bug
19974 | MALLOCBUFFERS | cannot allocate memory for all buffers | out of memory
19975 | ENQUEUEBUFFERINIT | enqueuing a buffer to the 'empty' queue failed | either queue is full or QID error
19976 | MARKQUEUED | cannot push to 'queued' queue | software bug
19977 | GETNEWBUFFEREMPTY | error accessing the 'empty' queue | software bug
19978 | GETNEWBUFFERTXED | error accessing the 'transmitted' queue | software bug
19979 | MARKEMPTY | cannot push to 'empty' queue | software bug
19980 | MARKTRANSMITTED | cannot push to 'transmitted' queue | software bug
19981 | MARKWARNING | cannot push to 'warning' queue | software bug
19982 | GETNEXTQUEUED | cannot access the 'queued' queue | software bug
19983 | GETNEXTWARNING | cannot access the 'warning' queue | software bug
19984 | MALLOCIMGBLOCK | cannot allocate 64 kBytes for image block buffer | out of resources?!?!
19985 | OPENIMGFILE | cannot open image file | open imagefile to continue from disk failed
19986 | IMAGELARGERTHANDEFINED | the file found on disk is larger than the control data indicates | software bug
19987 | CIS_OPENIMGFILE | cannot open the image file in correctImageSize | software bug
19988 | CIS_WRITEIMGFILE | cannot fill incomplete image to required size | check hardware (CF-Card)
19989 | DESTNOTAVAILABLE | destination is not available or returned error | Either the transmission failed due to a network problem or the target of an image transmission did not react or did reject the image for any reason. Check the destination. Image can be rerouted.
19990 | WRITEPIXELSTODISK | write error in imagefile | Meaning: writing the pixels to disk failed. Reason: CF-card defective or not enough memory left in partion e:/ Cure: replace CF-card.
19991 | WRITECTRLTODISK | write error in controlfile | Meaning: writing the controlfile to disk failed. Reason: CF-card defective or not enough memory left in partion e:/ Cure: replace CF-card.
19992 | NODISKSPACE | no enough space for spooler | Meaning: The available diskspace on e:/ is not sufficient for the spooler. Reason: Either too many files stored on e:/ or disk capacity decreased below an acceptable value. Cure: Install the newest software version with option format and restore previously saved backup data. If this does not help, exchange CF-card if the PMI board uses a CF-card, otherwise replace PMI board.
19993 | DESTNOTAVAILABLE1DEST | destination is not available or returned error | Either the transmission failed due to a network problem or the target of an image transmission did not react or did reject the image for any reason. Check the destination. No other destination available.
19994 | DEL_TIMEOUT | Timeout deleting image. | Meaning: Waiting for the end of a previous delete image request timed out. New request will not be executed. Cure: Try again later or delete the image folder manually and reset.
19995 | DEL_TIMEOUT2 | Timeout deleting image. | Meaning: Waiting for the end of a delete image request timed out. Request was cancelled. Cure: Try again later or delete the image folder manually and reset.
20480 | GET_PSDATA | getCPFData_ps failed | Retrieving the PS data from the CPF file failed
20481 | GET_APPLDATA | getCPFData_appl failed | Retrieving the APPL data from the CPF file failed
20482 | GET_NETDATA | getCPFData_net failed | Retrieving the NET data from the CPF file failed
20483 | GET_DIGDATA | getCPFData_dig failed | Retrieving the DIG data from the CPF file failed
20484 | ADD_DEST_TO_DS | AddDataToStore failed for DS_CPF_DESTCOORD | Adding the Destination Coordinates to the Data Store failed
20485 | ADD_STATIONNAME_TO_DS | AddDataToStore failed for DS_CPF_DIGCOORD | Adding the Station Names Coordinates to the Data Store failed
20486 | NOMEM_DEST | malloc returned NULL | No memory can be allocated for the structure holding all Destinations
20487 | GET_PRECDATA | getCPFData_prec_failed | Retrieving the PREC data from the CPF file failed
20488 | NOMEM_PREC | malloc returned NULL | No memory can be allocated for the structure holding all precheck destinations
20489 | MAINTENANCE_EXPECTED | Maintenance Interval will expire within a short time. Please contact service. | --
20490 | MAINTENANCE_REQUIRED | Maintenance Interval is expired. Please contact service as soon as possible. | --
20491 | GET_IDENDATA | getCPFData_iden failed | Retrieving the IDEN data from the CPF file failed
20492 | INFOCOUNTER_RECREATED | Infocounter data file recreated with default values. | Infocounter data file was recreated with default values, because original and backup data files were corrupted or lost.
20992 | GROUPNAME_SMALLER_THAN_ZERO | The group name is smaller than zero | --
20993 | GROUPNAME_TOO_BIG | The group name is too large | --
20994 | NOBUFFER | InitializeDataStore failed | Not enough memory could be allocated to hold the Data Store
20995 | NOT_ENOUGH_MEMORY | AddDataToStore failed | Not enough memory could be allocated
20996 | NOT_INITIALIZED | DataStore not initialized yet | Call 'InitializeDataStore' first.
20997 | TIMEOUT | lock already taken | A lock was taken and caused a waiting method to time out
21504 | General | General image processing error occurred | Meaning: A non-specific error occurred during image processing. Reason: Unknown. Cure: Unknown.
21505 | MFS_FsLength | Too few space for CRFileSystem. | Meaning: While trying to mount CRFileSystem, detected too few disk space. Reason: The flash memory on the photomultiplier unit is not existing or too small. Cure: Replace photomultiplier unit.
21506 | MFS_Mount | Could not mount CRFileSystem. | Meaning: Could not mount CRFileSystem. Reason: . Cure: Perform a scan line calibration.
21507 | PRP_NoCRImage | No valid CRImage given for preparing processing. | Meaning: The event to prepare processing contained no valid CRImage. Reason: Programming error. Either CRImage or one of its elements is a NULL pointer. Cure: Install another software version.
21508 | PRP_NoDoseCalPIP | Dose detection for PIP not calibrated. | Meaning: Dose detection for powder image plates not calibrated. Reason: Dose detection has not been calibrated at production or has been deleted. Cure: NIP may be used. For PIP, replace photomultiplier module.
21509 | PRP_NoDoseCalNIP | Dose detection for NIP not calibrated. | Meaning: Dose detection for needle image plates not calibrated. Reason: Dose detection has not been calibrated at production or has been deleted. Cure: PIP may be used. For NIP, replace photomultiplier module.
21514 | PRO_NoCRImage | No valid CRImage given for processing. | Meaning: The event to process an image contained no valid CRImage. Reason: Programming error. Either CRImage or one of its elements is a NULL pointer. Cure: Install another software version.
21515 | PRO_WriteRaw | Writing raw pixels to file failed. | Meaning: Writing raw pixels to file failed. Reason: Unknown. Cure: Ignore.
21524 | INI_NoDoseCal | Found no dose detection calibration at all. | Meaning: Found no dose detection calibration at all. Reason: Dose detection has not been calibrated at production or has been deleted. Cure: Replace photomultiplier module.
21525 | INI_NoCRTools | CRTools not started. | Meaning: At system start, the CRTools library was not found in memory. Reason: CRTools was not started. Cure: Install software again.
21526 | INI_IncompatibleCRTools | CRTools incompatible. | Meaning: The CRTools library is incompatible. Reason: The library CRTools does not fit to the rest of the application software. This is software configuration error and should never occur in the field. Cure: Try to install software again. If this doesn't help, install another software.
21554 | IMB_INI_BufferCreation | Creation of image buffer failed. | Meaning: Could not create a buffer for the image to scan. Reason: Memory for image buffer could not be allocated. Too few memory or problem with other parts of the software. Cure: Install the complete system again. If this fails, install another software version. Replace the CPU board.
21555 | IMB_PRP_TooManyColumns | Too many columns for preparing image buffer. | Meaning: At prepare operation, the number of coulumns and the block size have a wrong relation. Reason: Programming error. Cure: Install another software version.
21556 | IMB_GNW_Cancelled | Image transfer cancelled | Meaning: The image transfer has been cancelled. Reason: This may have many reasons. Cure: This is no error. Ignore it.
21557 | IMB_GNP_Cancelled | Image transfer cancelled | Meaning: The image transfer has been cancelled. Reason: This may have many reasons. Cure: This is no error. Ignore it.
21558 | IMB_GNP_Timeout | Image data timeout. | Meaning: Waiting for image data to process timed out. Reason: Could be programming error. Cure: Install another software version.
21559 | IMB_GNR_Cancelled | Image transfer cancelled | Meaning: The image transfer has been cancelled. Reason: This may have many reasons. Cure: This is no error. Ignore it.
21560 | IMB_GNR_Timeout | Image data timeout. | Meaning: Waiting for image data to process timed out. Reason: Could be programming error. Cure: Install another software version.
21561 | IMB_GAR_Cancelled | Image transfer cancelled | Meaning: The image transfer has been cancelled. Reason: This may have many reasons. Cure: This is no error. Ignore it.
21562 | IMB_GAR_Timeout | Image data timeout. | Meaning: Waiting for image data to process timed out. Reason: Could be programming error. Cure: Install another software version.
21563 | IMB_GAR_Occupied | All read blocks occupied. | Meaning: Could not get a new image block to read. All occupied. Reason: Could be programming error. Cure: Install another software version.
21564 | IMB_RAR_InvalidID | Tried to release an image block with an invalid ID. | Meaning: Tried to release an image block, which has not been reserved. Reason: Programming error. Cure: Ignore this error, if it does not cause any problem.
21565 | IMB_GNW_Overflow | Image buffer overflow | Meaning: The image buffer has not been read out fast enough, so that it is not ready for further image data. Scanning has been stopped. Reason: Image data was not stored or sent fast enough. Most probably network problem, but may also be too slow flash memory. Cure: - Check and correct the network connection. Available image data will be retransmitted automatically. Note that by switching off the device parts of the image data may be lost. - If the network and the PC installation are ok and the error occurs again, replace the digitizer CPU board. - At devices with CF-card try if replacing the flash memory by a faster one helps.
21566 | IMB_GAR_Released | Image data no longer available. | Meaning: Image data is no longer available. The memory location is reused in circular buffer. Reason: Could be programming error. Cure: Install another software version.
21567 | IMB_RBL_NoData | Attempt to release more data than available. | Meaning: Attempt to release more data than available. Corrected to actual value. Reason: Programming error. Cure: Only warning, should be harmless. In case of problems install another software version.
21604 | ISP_SignalRejected | The incoming signal was rejected in this state. | Meaning: Image splitter got a signal in a state where it cannot be handled. Reason: Programming error. Cure: In case of problems install another software version.
21605 | ISP_ImageQueueFull | The spooler image queue is full. | Meaning: Image queue of software component spooler is full. Reason: The spooler's capacity is exhausted, no image can be scanned at the moment. Cure: Wait until images are transmitted to the processing station. Be sure that the processing station is connected, configured correctly and working properly.
22016 | SYNCHRONIZING_FAILED | Synchronizing failed. | --
22528 | CRT_Base | General error occurred in CRTools library | Meaning: A non-specific error occurred in the CRTools library. Reason: This error should normally not occur. Its value is used as base for all other CRTools errors. Cure: Ignore this error.
22529 | PXC_InsufficientCalWidth | Calibration line does not cover the complete width of the image to be corrected. Parts of the image will not be corrected. | Meaning: Only parts of the image will be corrected. The rest might have banding artefacts. Reason: Calibration line does not cover the complete width of the image to be corrected. Cure: Calibrate with the biggest used cassette.
22530 | PXC_CalcCorrLines | Calculating the correction lines failed. Using constant correction line. | Meaning: The image is uncorrected. Reason: An exception occurred during preparation of the correction line. Might be software bug. Cure: Check the log file to find out more information about the error.
22531 | PXC_NoCalibrationLine | No calibration line loaded to correct the image. Using constant correction line. | Meaning: The image is uncorrected. Reason: Calibration line of this image plate type not available, invalid or could not be loaded. Cure: Calibrate the image plate type, which has generated the error.
22532 | PXC_InvalidIPType | An invalid IPType was given when clearing the calibration line. | --
22533 | PXC_ExtrapolateCorrLine | Extrapolating the correction lines failed. Using constant correction line. | Meaning: The image is uncorrected. Reason: An exception occurred during preparation of the correction line. Might be software bug. Cure: Check the log file to find out more information about the error.
22534 | PXC_UnsupportedSamplesOrBPP | Unsupported samples or bytes per pixel. | --
22535 | PXC_UnsupportedSamplesOrBPP2 | Unsupported samples or bytes per pixel. | --
22536 | PXC_UnsupportedColumns | Unsupported number of pixels per line. | Meaning: The requested number of pixels per line is negative or too small. Reason: Software bug. Cure: Install another software.
22537 | PXC_UnsupportedLines | Unsupported number of lines. | Meaning: The requested number lines is negative or too small. Reason: Software bug. Cure: Install another software.
22538 | LCL_CreateCalibrationLine | Creating a new CalibrationLine failed. | --
22539 | LCL_WriteToCRFile | Writing CalibrationLine to PMM failed. | --
22540 | DCL_CreateCalibrationLine | Creating a new dose CalibrationLine failed. | --
22541 | DCL_WriteToCRFile | Writing dose CalibrationLine to PMM failed. | --
22542 | PXC_LoadIPCalNullResult | PixelCorrector_LoadIPCal called with result null pointer. | Meaning: Null pointer in fuction call. Cannot execute. Reason: Software error. Cure: Install new software.
22543 | PXC_CorrectLinesParameter | PixelCorrector_CorrectLines called with invalid parameter. | Meaning: PixelCorrector_CorrectLines called with invalid parameter. Reason: Software error. Cure: Install new software.
22544 | PXC_EOIBlocksPending | There are still image blocks pending at end of image. Software error. | --
22545 | PXC_ExchangedStim | Exchanged optics module detected. | Meaning: Detected an exchanged optics module. Shading calibration not good enough any more. Reason: Replacement by spare part. Cure: Perform a shading calibration.
22546 | PXC_AcqSerNo | Wrong PMM serial number. | Meaning: Detected wrong photomultiplier module serial number in calibration line. Shading calibration not valid. Reason: PMM serial number was changed or wrong calibration line was manually copied. Cure: Perform a shading calibration.
22548 | CIM_UnknownImageFormat | Reading IP calibration image failed because of unknown image file format. | --
22549 | CIM_ImageNotLoaded | Reading IP calibration image failed. | --
22550 | CIM_ReadImageMemoryException | Exception at reading IP calibration image from ID tag. | --
22551 | CIM_ReadImageFileNameMismatch | Wrong IPSerialNumber in IP calibration file. | Meaning: The name of the IP calibration image is different than the IPSerialNumber, stored in the file. Reason: Probably manual renaming of the calibration file. Cure: Never use an IP calibration file for another image plate.
22552 | CIM_ReadImageFileException | Exception at reading IP calibration image from file. | --
22553 | CIM_WriteImageMemoryException | Exception at writing IP calibration image to memory for storing on ID tag. | --
22554 | CIM_WriteImageFileException | Exception at writing IP calibration image to file. | --
22555 | CIM_WriteNoImage | Writing IP calibration image failed because no image present. Software error. | --
22558 | CIM_RescaleNoImage | No image to rescale. | Meaning: There is no calibration image available, neither on the ID tag nor on the hard disk. Reason: This is a powder IP, where no IP calibration is necessary or the needle IP was not calibrated or has lost its calibration or IP correction was disabled for this scan. Cure: Return the image plate to production for calibration.
22559 | CIM_RescaleIPCalDotsPerMeter | The IP calibration image (TIFF file) contains wrong values for DotsPerMeter. | Meaning: Parameter error in the image plate calibration image. Reason: This could be prototype plate from DX-S development. But usually, this problem should not occur. Cure: Delete the calibration image from the hard disk and return the image plate to production for calibration.
22560 | CIM_RescaleIPCalFailed | Rescaling the given IP calibration image failed. | --
22561 | CIM_RescaleException | Exception at rescaling IP calibration image. | --
22568 | CIM_ShiftAddCutNoCal | No IP calibration image available. | Meaning: There is no calibration image available, neither on the ID tag nor on the hard disk. Reason: This is a powder IP, where no IP calibration is necessary or the needle IP was not calibrated or has lost its calibration. Cure: Return the image plate to production for calibration.
22569 | CIM_ShiftAddCutException | Exception at shifting / extending / cutting IP calibration image. | --
22570 | CIM_ShiftAddCutTopExtra | Too many lines on top of the calibration image had to be extrapolated. | Meaning: Too many lines on top of the calibration image had to be extrapolated. Possible image quality problem. Reason: The image plate was calibrated with completely different size settings than the current scan or the calibration image is for a different scan format. This can be caused by a duplicate definition of the image plate serial number. Cure: Return the image plate to production for new calibration and check of the serial number.
22571 | CIM_ShiftAddCutBottomExtra | Too many lines at the bottom of the calibration image had to be extrapolated. | Meaning: Too many lines at the bottom of the calibration image had to be extrapolated. Possible image quality problem. Reason: The image plate was calibrated with completely different size settings than the current scan or the calibration image is for a different scan format. This can be caused by a duplicate definition of the image plate serial number. Cure: Return the image plate to production for new calibration and check of the serial number.
22572 | CIM_ShiftAddCutLeftExtra | Too many columns at the left side of the calibration image had to be extrapolated. | Meaning: Too many columns at the left side of the calibration image had to be extrapolated. Possible image quality problem. Reason: The image plate was calibrated with completely different size settings than the current scan or the calibration image is for a different scan format. This can be caused by a duplicate definition of the image plate serial number. Cure: Return the image plate to production for new calibration and check of the serial number.
22573 | CIM_ShiftAddCutRightExtra | Too many columns at the right side of the calibration image had to be extrapolated. | Meaning: Too many columns at the right side of the calibration image had to be extrapolated. Possible image quality problem. Reason: The image plate was calibrated with completely different size settings than the current scan or the calibration image is for a different scan format. This can be caused by a duplicate definition of the image plate serial number. Cure: Return the image plate to production for new calibration and check of the serial number.
22574 | CIM_AddIPCalFile | Could not load IPCalFile. | --
22575 | CIM_AddInvalid | Invalid CalibrationImage object. | --
22576 | CIM_WriteMemoryReadFile | Could not read IPCalFile to write to memory. | --
22577 | CIM_WriteMemoryInvalid | Invalid CalibrationImage object. | --
22578 | CIM_WriteFileInvalid | Invalid CalibrationImage object. | --
22579 | CIM_WriteFileRescaledInvalid | Invalid CalibrationImage object. | --
22580 | CIM_RescaleAndFlipException | Error occurred while rescaling a calibration image. | --
22581 | CIM_WriteFileTrailerRead | Could not read IPCalFile trailer. | --
22582 | CIM_WriteFileTrailerWrite | Could not write IPCalFile trailer. | --
22583 | CIM_ReadFileTrailerInvalid | Invalid CalibrationImage object. | --
22584 | CIM_HeaderVersion | Wrong IPCalibration header version. | --
22585 | CIM_HeaderTruncated | IPCal header truncated. | --
22586 | CIM_DataTruncated | IPCalibration data truncated. | --
22587 | CIM_HeaderSize | Wrong IPCal header size. | --
22588 | CIM_HeaderType | Wrong IPCal header type. | --
22589 | CIM_IPCalHeaderVersion | Wrong IPCal header version. | --
22590 | CIM_IPCalHeaderTruncated | IPCalFile header truncated. | --
22591 | CIM_IPCalDataTruncated | Not enough IPCalFile data. | --
22592 | CIM_IPCalHeaderSize | Wrong IPCalFile header size. | --
22593 | CIM_Checksum | IPCalibrationData checksum error. | --
22594 | CIM_FileFormat | Unknown image file format. | --
22595 | CIM_SerialNumberMismatch | The image file name does not match the IP serial number. | --
22596 | CIM_WriteMNotValid | The IPCalibrationData to be written to tag is not valid. | Meaning: The IP calibration data to be written to the tag is not valid. Reason: Probably software error. Cure: Install new software.
22597 | CIM_WriteMTooBig | The IPCalData to be written to tag is too big. | Meaning: Detected that the given IP calibration data does not fit onto the tag. Reason: Probably software error. Cure: Install new software.
22598 | CIM_WriteMNotValid2 | The IPCalibrationData to be written to tag is not valid. | --
22599 | CIM_WriteNotValid | The IPCalFile data to be written to file is not valid. | --
22600 | CIM_NoFile | Could not open CalibrationImage file. | Meaning: There is no calibration image available, neither on the ID tag nor on the hard disk. Reason: . Cure: Return the image plate to production for new calibration and check of the serial number.
22601 | CIM_IPCalDataVersion | Wrong IPCalData version. | --
22602 | CIM_IPCalDataVersion2 | Wrong IPCalData version. | --
22603 | CIM_IPCalDataType | Wrong IPCalData type. | --
22604 | CIM_IPCalDataTrailerLen | IPCalFile trailer longer than data. | --
22605 | CIM_IPCalFileTooSmall | IPCalFile too small. | --
22606 | CIM_IPCalFileChecksum | IPCalFile checksum error. | --
22607 | CIM_WriteThumbnailInvalid | Invalid CalibrationImage object. | --
22608 | CIM_WriteThumbnailNoData | CalibrationImage object contains no image data. | --
22609 | CIM_WriteThumbnailRescale | Rescaling the CalibrationImage failed. | --
22610 | CIM_WriteThumbnailConvert | Converting the CalibrationImage to 8 bits failed. | --
22611 | CIM_AddNewAllocImage | Could not allocate new image, when adding a new (empty) IP calibration image. | --
22612 | CIM_AddNewIPCalFile | Adding a new (empty) IP calibration image failed. | --
22613 | CIM_AddNewInvalid | Invalid CalibrationImage object. | --
22614 | CIM_CalibrateInvalid | Invalid CalibrationImage object. | --
22615 | CIM_CalibrateNoData | CalibrationImage object contains no image data. | --
22616 | CIM_NormalizeInvalid | Invalid image for normalization. | --
22617 | CIM_Normalize0 | Could not normalize image. All pixels are 0. | --
22618 | CIM_NoBinaryData | No data given to create a CalibrationImage from. | --
22619 | CIM_NoSingleBlock | ImageArea does not consist of one single block. | Meaning: The ImageArea for line artefact correction does not consist of one single block. No correction done. Reason: Software error. Cure: No cure. The image is ok, besides the fact that a possible line artefact is not corrected.
22628 | CLN_CreateType | Invalid CalibrationLine file type. | --
22629 | CLN_CreatePixelWidth | Invalid pixelWidth in CalibrationLine file. | --
22630 | CLN_CreateNLines | Invalid number of calibration lines in CalibrationLine file. | --
22631 | CLN_CreateNColumns | Invalid number of columns in CalibrationLine file. | --
22632 | CLN_ROILeft | Invalid ROILeft in CalibrationLine file. | --
22633 | CLN_CreateBitsPP | Unsupported number of bits per pixel in CalibrationLine file. | --
22634 | CLN_CreateOpenFile | Opening the CalibrationLine file failed. | --
22635 | CLN_CreateOpenCRFile | Opening the CalibrationLine file in CRFileSystem failed. | --
22636 | CLN_DeserializeInvalid | Invalid CalibrationLine object. | --
22637 | CLN_DeserializeLength | CalibrationLine file too small. | --
22638 | CLN_DeserializeFileType | Invalid CalibrationLine file type. | --
22639 | CLN_DeserializeHeaderLength | Invalid CalibrationLine file header length. | --
22640 | CLN_DeserializeHeaderLength2 | Invalid CalibrationLine file header length. | --
22641 | CLN_DeserializeHeaderLength1 | Invalid CalibrationLine file header length. | --
22642 | CLN_DeserializeVersion | Invalid version of CalibrationLine file. | --
22643 | CLN_DeserializeLength2 | CalibrationLine file too small. | --
22644 | CLN_DeserializePixelWidth1 | Invalid pixelWidth in CalibrationLine file. | --
22645 | CLN_DeserializePixelWidth | Invalid pixelWidth in CalibrationLine file. | --
22646 | CLN_DeserializeNColumns | Invalid number of columns in CalibrationLine file. | --
22647 | CLN_DeserializeBitsPP | Unsupported number of bits per pixel in CalibrationLine file. | --
22648 | CLN_DeserializeLength3 | CalibrationLine file too small. | --
22649 | CLN_DeserializeNLines | Invalid number of calibration lines in CalibrationLine file. | --
22650 | CLN_DeserializeROILeft | Invalid ROILeft in calibration file. | --
22651 | CLN_DeserializeLength4 | Too few data to deserialize calibration line. | --
22652 | CLN_SerializeInvalid | CalibrationLine_Serialize: Empty calibration line. | --
22653 | CLN_SerializeHeaderBytes | CalibrationLine_Serialize: Wrong HeaderBytes (software bug). | --
22654 | CLN_NormalizeInvalid | CalibrationLine_Normalize: Empty calibration line. | --
22655 | CLN_ClipInvalid | Clip: Empty calibration line. | --
22656 | CLN_MaximumSlopeInvalid | CalibrationLine_MaximumSlope: Empty calibration line. | --
22657 | CLN_MaximumSlopeNegative | CalibrationLine_MaximumSlope: Unexpected negative value of maxDifference. | --
22658 | CLN_InterpolateInvalid | CalibrationLine_Interpolate: Empty calibration line. | --
22659 | CLN_InterpolatePixelWidth | CalibrationLine_Interpolate: Invalid new pixel width. | --
22660 | CLN_RescaleNNMatch | CalibrationLine_RescaleNearestNeighbor: Length of input line and number of sections do not match. | --
22661 | CLN_RescaleCubicMatch | CalibrationLine_RescaleCubic: Length of input line and number of sections do not match. | --
22662 | CLN_ExtrapolateInvalid | CalibrationLine_Extrapolate: Empty calibration line. | --
22663 | CLN_ExtrapolateCut | CalibrationLine_Extrapolate: Cannot cut more left columns than available. | --
22664 | CLN_ExtrapolateExtend | CalibrationLine_Extrapolate: Cannot extend beyond begin of line. | --
22665 | CLN_ExtrapolateColumns | CalibrationLine_Extrapolate: Invalid number of new columns. | --
22666 | CLN_MultiplyInvalid | CalibrationLine_Multiply: Empty calibration line. | --
22667 | CLN_MultiplyLinesInvalid | CalibrationLine_Multiply: Empty calibration line. | --
22668 | CLN_MultiplyLinesInvalid2 | CalibrationLine_Multiply: Empty calibration line. | --
22669 | CLN_MultiplyLinesBitsPP | CalibrationLine_Multiply: PixelBits do not match. | --
22670 | CLN_MultiplyLinesNCalLines | CalibrationLine_Multiply: NumberOfCalLines do not match. | --
22671 | CLN_DivideInvalid | CalibrationLine_Divide: Empty calibration line. | --
22672 | CLN_DivideZero | CalibrationLine_Divide: Division by zero. | --
22673 | CLN_SubtractInvalid | CalibrationLine_Subtract: Empty minuend line. | --
22674 | CLN_SubtractInvalid2 | CalibrationLine_Subtract: Empty subtrahend line. | --
22675 | CLN_ExLineInvalid | CalibrationLine_ExtrapolateLine: Line to be extrapolated is NULL! | --
22676 | CLN_ExLineStart | CalibrationLine_ExtrapolateLine: startIndex not valid for line. | --
22677 | CLN_ExLineLength | CalibrationLine_ExtrapolateLine: lineLength not valid for line. | --
22678 | CLN_ExLineFirstvalid | CalibrationLine_ExtrapolateLine: firstValid is out of range. | --
22679 | CLN_ExLineLastValid | CalibrationLine_ExtrapolateLine: lastValid is out of range. | --
22680 | CLN_MinMaxScanValInvalidLine | CalibrationLine_CalculateMinMaxScanVal: Empty calibration line. | --
22681 | CLN_SplineDivByZero | Bad input to routine splint leads to division by 0. | --
22682 | CLN_SetCalLinesLength | Could not set CalLines due to wrong length. | --
22683 | CLN_SetCalLinesMalloc | CalibrationLine_set_CalLines failed. malloc error. | --
22684 | CLN_ConsistencyInvalid | CalibrationLine_CheckConsistency: Empty calibration line. | --
22685 | CLN_ConsistencyNotNormalized | CalibrationLine_CheckConsistency: Gain line is not normalized. | --
22686 | CLN_ConsistencyRange | CalibrationLine_CheckConsistency: Gain line contains too low pixel values. | --
22687 | CLN_Checksum | CalibrationLine checksum error. | Meaning: There was a checksum error in a read calibration line. Reason: Transmission problem from photomultiplier module (PMM) to CPU board (PMI). Cure: The software automatically retries to read the calibration line. If the problem persists, perform a shading calibration.
22688 | CLN_DenormalizeInvalid | CalibrationLine_Denormalize: Empty calibration line. | --
22689 | CLN_DenormalizeNotNormalized | CalibrationLine_Denormalize: Calibration line is not normalized. | --
22708 | MED_USNoData | SlidingMedian_UShort: No data given to calculate median. | --
22709 | MED_USRangeTooSmall | SlidingMedian_UShort: The median range is too small. | --
22710 | MED_USRangeTooBig | SlidingMedian_UShort: The median range is bigger than the data length. | --
22711 | MED_FNoData | SlidingMedian_Float: No data given to calculate median. | --
22712 | MED_FRangeTooSmall | SlidingMedian_Float: The median range is too small. | --
22713 | MED_FRangeTooBig | SlidingMedian_Float: The median range is bigger than the data length. | --
22714 | ART_DetNoData | LineArtefact_Detect: No image data given. | --
22715 | ART_DetAreaSize | LineArtefact_Detect: Columns or line is less or equal 0 | --
22716 | ART_CorrNoData | LineArtefact_Correct: No image data given. | --
22717 | ART_CorrAreaSize | LineArtefact_Correct: Columns or line is less or equal 0 | --
22718 | RTC_QFull | Retrace correction block queue is full. Software error. | --
22719 | RTC_NotSuccessiveBlocks | Retrace correction got not successive blocks. Software error. | --
22720 | RTC_BlockTooSmall | First image block for retrace correction is too small. Software error. | --
22721 | RTC_NoInLine1 | Retrace correction did not find correct input line. Software error. | --
22722 | RTC_NoInLine2 | Retrace correction did not find correct input line. Software error. | --
22723 | RTC_NoInLine3 | Retrace correction did not find correct input line. Software error. | --
22724 | RTC_NoOutLine | Retrace correction did not find correct output line. Software error. | --
22728 | IMA_AddLines | Could not add lines to ImageArea. | --
22729 | GRA_Allocate | Could not allocate graph image. | --
22730 | GRA_Rescale | Could not rescale graph image. | --
22731 | GRA_Save | Could not save graph image. | --
22748 | SYN_CRFSNotMounted | SyncCalLines: CRFileSystem not mounted. | --
22749 | SYN_CRFSDir | SyncCalLines: Could not get CRFileSystem directory. | --
22750 | SYN_CRFSOpen | SyncCalLines: Could not open calibration file from CRFileSystem. | --
22751 | SYN_CRFSNoData | SyncCalLines: Calibration file on CRFileSystem contains no data. | --
22752 | SYN_Malloc | SyncCalLines: Could not allocate buffer for CRFile. | --
22753 | SYN_FOpen | SyncCalLines: Could not open calibration file for writing. | --
22754 | SYN_Dir | SyncCalLines: Could not get directory. | --
22755 | SYN_CreateCL | SyncCalLines: Could not create CalibrationLine from file. | --
22756 | SYN_CRFSWrite | SyncCalLines: Could not write CalibrationLine to CRFileSystem. | --
22757 | SYN_ExchangedAcq | SyncCalLines: Exchanged PMM detected. | Meaning: Detected an exchanged photomultiplier module. Shading calibration not good enough any more. Reason: Replacement by spare part. Cure: Perform a shading calibration.
22758 | SYN_ExchangedStim | SyncCalLines: Exchanged optics module detected. | Meaning: Detected an exchanged optics module. Shading calibration not good enough any more. Reason: Replacement by spare part. Cure: Perform a shading calibration.
22759 | SYN_AcqSerNo | SyncCalLines: Wrong PMM serial number. | Meaning: Detected wrong photomultiplier module serial number in calibration line. Shading calibration not valid. Reason: PMM serial number was changed or wrong calibration line was manually copied. Cure: Perform a shading calibration.
22760 | SYN_GetLastWriteTime | SyncCalLines: Could not get the CalibrationLine write time (disk). | --
22761 | SYN_CRFSSetLastWriteTime | SyncCalLines: Could not update the CalibrationLine write time in CRFileSystem. | --
22762 | SYN_FWrite | SyncCalLines: Could not write CalibrationLine to disk. | --
23040 | SetError | Setting an error failed. | --
23041 | RemoveError | Removing an error failed. | --
23042 | StartDiagnoseCycle | Start Diagnose Cycle | --
23043 | StartEnduranceRunCycle | Start Endurancerun Cycles | --
23044 | StartShadingCalibration | Start shading clibration | --
23045 | MoveIPCarrierToMaintenance | Device ready to move IP Carrier to maintenance position | --
23046 | MoveIPCarrierToPark | Device ready to move IP Carrier to park position | --
23049 | StartStallCalibration | Start Stall Calibration | --
23050 | NO_PMI_COUNTERS | No counters on PMI board | --
23051 | NO_SUCH_MODULE | Device has no such module | --
23052 | PREPSHIP_CANNOT_READ_DEVICE_ELABEL | Prepare shipment cannot read device eLabel | --
23053 | PREPSHIP_CANNOT_PROTECT_DEVICE_ELABEL | Prepare shipment cannot protect device eLabel | --
23054 | PREPSHIP_CANNOT_PROTECT_OPTICS_MODULE | Prepare shipment cannot protect optics module | --
23055 | PREPSHIP_CANNOT_PROTECT_PMM | Prepare shipment cannot protect PMM | --
23056 | PREPSHIP_CANNOT_CREATE_PRODINFO_FILE | Prepare shipment cannot create the product info file | --
23057 | PREPSHIP_CANNOT_ENCRYPT_TIME | Prepare shipment cannot encrypt the digitizer time | --
23058 | PREPSHIP_CANNOT_READ_SERIAL_NUMBER | Prepare shipment cannot read the product serial number from the device eLabel | --
23552 | START_SERVER_FAILED | Start Server failed | Starting of DirectID Server failed
24064 | WRONG_UIC_PROTVER | Protocol version of UIC is wrong | Meaning: Incompatible SW components (protocol version of UserInterface UIC is wrong). Reason: SW error. Cure: Install new SW version on all boards; inform R&amp;D.
24065 | NO_IP_PRESENT_AFTER_SCAN | No Ip at Sensor present after scan movement | Meaning: At the Ip Sensor Light Barrier no Ip is detected after the scan movement. Reason: Handling of the Ip not successful. Cure: 1) Restart Digitizer. 2) Open Digitzer and remove Ip.
24066 | SENSOR_DRAWER_HOME_NOT_CONNECTED | The home reference sensor of the drawer drive is not connected. | Meaning: The home reference sensor of the drawer drive is not connected. Reason: -- Cure: Check the sensor and the corresponding cabling. Replace the defective parts.
24067 | IP_SENSOR_NOT_CONNECTED | The ip sensor of the transport unit is not connected. | Meaning: The ip sensor of the transport unit is not connected. Reason: -- Cure: Check the sensor and the corresponding cabling. Replace the defective parts.
24068 | SENSOR_CASSETTE_DETECTION_NOT_CONNECTED | The cassette detection sensor of the cassette input slot is not connected. | Meaning: The cassette detection sensor of the cassette input slot is not connected. Reason: -- Cure: Check the sensor and the corresponding cabling. Replace the defective parts.
24069 | DRAWER_HANDOVER_IP_SEQUENCE_FAILURE | Drawer handoverIp sequence failure. | Meaning: The Drawer handoverIp sequence does not perform in expected order. Reason: Movement of the IP jammed at the shifting into the scanner. Drawer reference Light Barrier not reached as expected. Cure: Remove jammed IP and restart.
24070 | IP_PRESENT_AFTER_RETURN_IP_SSCAN_AT_CLEANUP | Ip at Sensor present after slowscan movement returnIp at CleanUp | Meaning: At the Ip Sensor Light Barrier an Ip is detected after the slowscan movement for returnIp at CleanUp. Reason: Handling of the Ip not successful. Cure: 1) Restart Digitizer. 2) Open Digitzer and remove Ip.
24071 | IP_PRESENT_AFTER_RETURN_IP_SSCAN | Ip at Sensor present after slowscan movement returnIp | Meaning: At the Ip Sensor Light Barrier an Ip is detected after the slowscan movement for returnIp. Reason: Handling of the Ip not successful. Cure: 1) Restart Digitizer. 2) Open Digitzer and remove Ip.
24072 | LOST_IP_PRESENT_AT_FEED_IP_BEGIN_FAILURE | Ip at Sensor present at feedIp movement begin. | Meaning: At the Ip Sensor Light Barrier an Ip is detected at movement for feedIp. Reason: Handling of the Ip not successful to bring back the Ip into the cassette. Cure: 1) Restart Digitizer. 2) Open Digitzer and remove Ip.
24073 | SENSOR_CABLE_DRAWER_LB_AND_CASS_DET_NOT_CONNECTED | The home reference sensor of the drawer drive and the cassette detection sensor is not connected. | Meaning: The home reference sensor of the drawer drive and the cassette detection sensor is not connected. The cable for both sensors is disconnected at the integrated plug. Reason: -- Cure: Check the plug for disconnection. Check the sensors and the corresponding cabling. Replace the defective parts.
24074 | CASS_RELEASED_DURING_SCAN | Cassette released during cycle or during clean-up | Meaning: Cassette sensor opened during scan or erasure cycle or during clean-up. Reason: User pressed the cassette release button during a running cycle or during the clean-up. This is not allowed, which is also described in the user manual. Or cassette detection mechanism is defective. Cure: 1) Clamp the cassette again by pushing it towards the digitizer and restart Digitizer. Check if image plate is scratched. 2) Check 'Cassette Detection Light Sensor' and cabling. 3) Check cassette unit mechanics. 4) Replace 'Power Board'.
24075 | IP_PRESENT_AFTER_ERASURE | IP at Sensor present at end of erasure | Meaning: At the IP Sensor Light Barrier an image plate (IP) is detected at end of erasure. Reason: Handling of the IP not successful, IP was not completely in the cassette when the erasure was finished. The digitizer moved the IP automatically back into the cassette by executing a rescue cycle, but the IP is not completely erased. Cure: Image plate is not erased! Remove the cassette and press the confirm button. Do not use the cassette again before an inspection. Refer to the detailed description in the user manual.
24164 | WARN_CLEANUP_BEGIN | CleanUp initiated. | The CleanUp prodedure is imposed to the handling
24165 | WARN_CLEANUP_FIN_FAILED | CleanUp finished with failure. | The CleanUp prodedure finished with failure.
24166 | WARN_CLEANUP_FIN_OK_IP_FOUND | CleanUp terminates successful with an IP detected. | The CleanUp prodedure terminates successful with an IP detected during the movement to the tray. The Ip is transported back to the cassette already.
24167 | WARN_CLEANUP_FIN_OK_IP_NOT_FOUND | CleanUp terminates successful without an IP detected. | The CleanUp prodedure terminates successful without an IP detected at the movement to the tray.
24583 | MOTOR_UNPLUGGED | Drawer Selftest UNPLUGGED Motor failure. | Meaning: Selftesting the motor results in a detection of the Unplugged failure. Reason: Current at all the electrical motor windings not sufficient. Cure: Retry, restart, check cabling, motor and Power Board.
24584 | MOTOR_DEFECTIVE | Drawer Selftest DEFECTIVE Motor failure. | Meaning: Selftesting the motor results in a detection of the Defective Motor failure. Reason: Current at one of the electrical motor windings not sufficient. Wire broken. Cure: Retry, restart, check cabling, motor and Power Board.
24585 | MOTOR_DRIVER | Drawer Selftest DRIVER failure. | Meaning: Selftesting the motor results in a detection of the Driver failure. Reason: Current at one of the electrical motor windings too high. Wires shorted. Motor has short circuit. Motor Current Driver circuit at Power Board defective. Cure: Retry, restart, check cabling, motor and Power Board.
24586 | MOTOR_SELFTEST | Drawer Selftest returns general SELFTEST failure. | Meaning: Selftesting the motor results in a detection of a unspecific general failure. Reason: Current at one of the electrical motor windings too high. Wires shorted. Motor has short circuit. Motor Current Driver circuit at Power Board defective. Also possible is: Current at one of the electrical motor windings not sufficient. Wire broken. Also possible is: Communication signals to Motor Driver have failure or internal Power Board function is broken. Cure: Retry, restart, check cabling, motor and Power Board.
24587 | FIND_REF_FAILURE | Procedure findRef of DRAWER terminated with general failure. | Meaning: Find Reference movement of the Dawer finished not successful.
24588 | FIND_REF_FAILED_AS_LB_NOT_REACHED | Procedure findRef of DRAWER terminated with Lb not Reached failure. | Find Reference of the Dawer finished not successful as the move into the Light Barrier terminated unexpected by Step Limitation. The sensor LB did not react as expected.
24589 | FIND_REF_FAILED_STEP_DEVIATION | Procedure findRef of DRAWER terminated with Step Deviation failure. | Find Reference of the Dawer finished not successful as the moves terminated unexpected by Step Deviation over limit. The movement forward and backward did not match as expected.
24590 | FIND_REF_FAILED_AS_LB_NO_PULLOUT | Procedure findRef of DRAWER terminated with Lb no pullout failure. | Find Reference of the Dawer finished not successful as the move out of the Light Barrier terminated unexpected. The sensor LB did not react as expected for pullout movement.
24591 | HANDOVER_IP_FAILED_WRONG_STARTPOS | Procedure handoverIp of DRAWER terminated with Lb not free - wrong startpos failure. | HandoverIp of the Dawer finished not successful as the move into the Light Barrier terminated unexpected as the startposition is wrong. The sensor LB did not react as expected for handoverIp movement. The sensor indicates an already covered LB.
24592 | HANDOVER_IP_FAILED_LB_NOT_REACHED | Procedure handoverIp of DRAWER terminated with Lb not reached failure. | HandoverIp of the Dawer finished not successful as the move into the Light Barrier terminated unexpected by Step Limitation. The sensor LB did not react as expected.
24593 | HANDOVER_IP_FAILED_IDLE_POS_NOT_REACHED | Procedure handoverIp of DRAWER terminated with idle pos not reached failure. | HandoverIp of the Dawer finished not successful as the move into the Idle Pos terminated unexpected.
24594 | RELEASE_IP_FAILED_WRONG_STARTPOS | Procedure releaseIp of DRAWER terminated with Lb not free - wrong startpos failure. | ReleaseIp of the Dawer finished not successful as the move into the Light Barrier terminated unexpected as the startposition is wrong. The sensor LB did not react as expected for releaseIp movement. The sensor indicates an already covered LB.
24595 | RELEASE_IP_FAILED_LB_NOT_REACHED | Procedure releaseIp of DRAWER terminated with Lb not reached failure. | ReleaseIp of the Dawer finished not successful as the move into the Light Barrier terminated unexpected by Step Limitation. The sensor LB did not react as expected.
24596 | RELEASE_IP_FAILED_IDLE_POS_NOT_REACHED | Procedure releaseIp of DRAWER terminated with idle pos not reached failure. | ReleaseIp of the Dawer finished not successful as the move into the Idle Pos terminated unexpected.
24597 | RELEASE_IP_FAILED_ABS_POS_DEVIATION | Procedure releaseIp of DRAWER terminated with absolute positioning deviation exceeded failure. | ReleaseIp of the Dawer finished not successful as the move into the Idle Pos terminated unexpected. The absolute positioning deviation exceeded the defined range. Possible situation is that the current or previous movement lost steps as an jam happened.
24598 | HANDOVER_IP_FAILED_ABS_POS_DEVIATION | Procedure handoverIp of DRAWER terminated with absolute positioning deviation exceeded failure. | HandoverIp of the Dawer finished not successful as the move into the Idle Pos terminated unexpected. The absolute positioning deviation exceeded the defined range. Possible situation is that the current or previous movement lost steps as an jam happened.
24599 | NO_REACTION_TIMEOUT_FAILURE | Procedure for movement of DRAWER terminated with general timeout failure. | Procedure for movement of DRAWER terminated with general timeout failure. No Reaction from motor movement. Movement finished not successful in expected time period.
25095 | RETURN_IP_FAILURE | Return Ip problem due to start problem or blocking failure. | Meaning: At the Return Ip Movement a problem is indicated. Reason: The start position is problematic or blocking is present. Oder and Sequence of Activities wrong. Cure: Remove IP and Cassette manually, Retry, Restart.
25096 | MOTOR_UNPLUGGED | SlowScan Selftest UNPLUGGED Motor failure. | Meaning: Selftesting the motor results in a detection of the Unplugged failure. Reason: Current at all the electrical motor windings not sufficient. Cure: Retry, restart, check cabling, motor and Power Board.
25097 | MOTOR_DEFECTIVE | SlowScan Selftest DEFECTIVE Motor failure. | Meaning: Selftesting the motor results in a detection of the Defective Motor failure. Reason: Current at one of the electrical motor windings not sufficient. Wire broken. Cure: Retry, restart, check cabling, motor and Power Board.
25098 | MOTOR_DRIVER | SlowScan Selftest DRIVER failure. | Meaning: Selftesting the motor results in a detection of the Driver failure. Reason: Current at one of the electrical motor windings too high. Wires shorted. Motor has short circuit. Motor Current Driver circuit at Power Board defective. Cure: Retry, restart, check cabling, motor and Power Board.
25099 | MOTOR_SELFTEST | SlowScan Selftest returns general SELFTEST failure. | Meaning: Selftesting the motor results in a detection of a unspecific general failure. Reason: Current at one of the electrical motor windings too high. Wires shorted. Motor has short circuit. Motor Current Driver circuit at Power Board defective. Also possible is: Current at one of the electrical motor windings not sufficient. Wire broken. Also possible is: Communication signals to Motor Driver have failure or internal Power Board function is broken. Cure: Retry, restart, check cabling, motor and Power Board.
25100 | NO_REACTION_TIMEOUT_FAILURE | SlowScan Movement returns general TIMEOUT failure. | A movement of SlowScan terminated with general Timeout failure. No Reaction from motor movement HW-control. Movement finished not successful in expected time period.
25101 | RETURN_IP_SYNC_DISTANCE_FAILURE | Warning Message as Return Ip move received invalid syncDistance. | Return Ip received invalid syncDistance parameter for Drawer movement. Use Default. Warning Message.
25102 | CALLBACK_POINTER_ZERO | The call-back pointer invalid. Obsolete as callbacks replaced by external port functions. | The call-back pointer has the value zero and therefore no response could be sent back from the corresponding passive class. Obsolete as callbacks replaced by external port functions.
25103 | WRONG_STATE_LOW_LEVEL_STATE_MACHINE | The received event cannot be handled in the actual state of the low level state machine of the passive class SlowScan. | Meaning: The received event cannot be handled in the actual state of the low level state machine of the passive class SlowScan. Reason: A wrong state or has been set. Cure: Check the corresponding SW.
25104 | SPEED_FACTOR_OUT_OF_REQUIRED_RANGE | The slowscan speed factor is out of the required range. | Meaning: The slowscan speed factor is out of the required range. Reason: The corresponding value was not correctly transferred or the value was not correctly calculated. Cure: Check the corresponding value and the SW-procedures.
25600 | FUSE_DEFECT_RED_LED | Erasure Unit reports a supply voltage error. | Meaning: The fuse for the supply voltage to the Erasure Unit is broken at the green LED line. Possible Reason: Short Circuit at the green LED line current path.
25601 | FUSE_DEFECT_GREEN_LED | Erasure Unit reports a supply voltage error. | Meaning: The fuse for the supply voltage to the Erasure Unit is broken at the red LED line. Possible Reason: Short Circuit at the red LED line current path.
25602 | MISSING_CURRENT_RED_LED | Erasure Unit reports LED operating current error. | Meaning: Erasure Unit reports that LED operating current is not measured although it is switched on. The Erasure Unit could not be switched on successfully. Red LED line current is not OK. Possible Reasons: Red LED line current path broken or disconnected. Overtemperature protection at plug P2 is disconnected or has broken wires. The Erasure Unit needs to cool down if too hot.
25603 | MISSING_CURRENT_GREEN_LED | Erasure Unit reports LED operating current error. | Meaning: Erasure Unit reports that LED operating current is not measured although it is switched on. The Erasure Unit could not be switched on successfully. Green LED line current is not OK. Possible Reasons: Green LED line current path broken or disconnected. Overtemperature protection at plug P2 is disconnected or has broken wires. The Erasure Unit needs to cool down if too hot.
25604 | HIGH_OFF_CURRENT_RED_LED | Erasure Unit reports LED quiescent current error. | Meaning: Erasure Unit reports that LED operating current is measured although it is switched off. The Erasure Unit could not be switched off successfully. Red LED line current is not Off. Possible Reason: The 'Power Board' electronics has a defect at the driver or the sensoric for the Erasure Unit.
25605 | HIGH_OFF_CURRENT_GREEN_LED | Erasure Unit reports LED quiescent current error. | Meaning: Erasure Unit reports that LED operating current is measured although it is switched off. The Erasure Unit could not be switched off successfully. Green LED line current is not Off. Reason: The 'Power Board' electronics has a defect at the driver or the sensoric for the Erasure Unit.
25606 | ERASURE_MODULE_DISCONNECTED | Erasure Unit reports LED operating current error. Disconnection possible | Meaning: Erasure Unit reports that LED operating current is not measured although it is switched on. The Erasure Unit could not be switched on successfully. Green and Red LED line currents are not OK. Possible Reasons: Green and Red LED line current paths broken or disconnected. Overtemperature protection at plug P2 is disconnected or has broken wires. The Erasure Unit needs to cool down if too hot.
25607 | FUSE_DEFECT_RED_LED_SELFTEST | Erasure Unit reports a supply voltage error. | Meaning: The fuse for the supply voltage to the Erasure Unit is broken at the green LED line. Possible Reason: Short Circuit at the green LED line current path.
25608 | FUSE_DEFECT_GREEN_LED_SELFTEST | Erasure Unit reports a supply voltage error. | Meaning: The fuse for the supply voltage to the Erasure Unit is broken at the red LED line. Possible Reason: Short Circuit at the red LED line current path.
25609 | MISSING_CURRENT_RED_LED_SELFTEST | Erasure Unit reports LED operating current error. | Meaning: Erasure Unit reports that LED operating current is not measured although it is switched on. The Erasure Unit could not be switched on successfully. Red LED line current is not OK. Possible Reasons: Red LED line current path broken or disconnected. Overtemperature protection at plug P2 is disconnected or has broken wires. The Erasure Unit needs to cool down if too hot.
25610 | MISSING_CURRENT_GREEN_LED_SELFTEST | Erasure Unit reports LED operating current error. | Meaning: Erasure Unit reports that LED operating current is not measured although it is switched on. The Erasure Unit could not be switched on successfully. Green LED line current is not OK. Possible Reasons: Green LED line current path broken or disconnected. Overtemperature protection at plug P2 is disconnected or has broken wires. The Erasure Unit needs to cool down if too hot.
25611 | HIGH_OFF_CURRENT_RED_LED_SELFTEST | Erasure Unit reports LED quiescent current error. | Meaning: Erasure Unit reports that LED operating current is measured although it is switched off. The Erasure Unit could not be switched off successfully. Red LED line current is not Off. Possible Reason: The 'Power Board' electronics has a defect at the driver or the sensoric for the Erasure Unit.
25612 | HIGH_OFF_CURRENT_GREEN_LED_SELFTEST | Erasure Unit reports LED quiescent current error. | Meaning: Erasure Unit reports that LED operating current is measured although it is switched off. The Erasure Unit could not be switched off successfully. Green LED line current is not Off. Reason: The 'Power Board' electronics has a defect at the driver or the sensoric for the Erasure Unit.
25613 | ERASURE_MODULE_DISCONNECTED_SELFTEST | Erasure Unit reports LED operating current error. Disconnection possible | Meaning: Erasure Unit reports that LED operating current is not measured although it is switched on. The Erasure Unit could not be switched on successfully. Green and Red LED line currents are not OK. Possible Reasons: Green and Red LED line current paths broken or disconnected. Overtemperature protection at plug P2 is disconnected or has broken wires. The Erasure Unit needs to cool down if too hot.
26112 | SW_VERSION_MISMATCH | The software of the digitizer and the ErrorViewer on PC do not match. | Meaning: The software of the digitizer and the ErrorViewer on NX workstation do not match. The ErrorViewer is not able to display messages of the digitizer. Reason: 1. An old version of ErrorViewer is installed on NX workstation. Cure: Please install the ErrorViewer on NX wordkstation released for actual version of digitizer software.
26113 | CREATE_SENDER_SOCKET | Could not create sender socket. | Meaning: Internal software error connecting to NX workstation. Reason: TCP/IP communication problem with NX workstation Cure: Please restart the digitizer. If the problem occurs after restart again, please call service.
26114 | CONNECT_FAILED | Could not connect to NX workstation. | Meaning: Connection of ErrorViewer to NX failed Reason: 1. The NX workstation is not available in the network. 2. The NX workstation is not started or user is not logged in at NX workstation 3. The ErrorViewer on NX workstation is not started. 4. The ErrorViewer on NX workstation is started but it has a problem and cannot connect and receive messages 5. ErrorViewer version on NX workstation is not supported 6. The digitizer is configured for another NX workstation than the one actually used. Cure: 1. Check network cabling connection to NX workstation and digitizer. 2. Start NX workstation and login. 3. Start the ErrorViewer on NX workstation. 4. Exit the ErrorViewer and restart it. 5. Install an ErrorViewer with version 2.0 or higher and start the ErrorViewer on NX workstation. 5. Reconfigure NX workstation and digitizer.
26115 | TIMEOUT_SEND | No response from NX workstation | Meaning:NX workstation does not respond. Reason: 1. The ErrorViewer on NX workstation is started but it has a problem and cannot connect and receive messages 2. ErrorViewer version on NX workstation is not supported Cure: 1. Exit the ErrorViewer and restart it. 2. Install an ErrorViewer with version 2.0 or higher
26117 | ACK_NOT_SYNC | Acknoledge asynchron | Meaning: Communication problem between between digitizer and Erroviewer on NX wordkstation Reason: Communication between digitizer and Erroviewer on NX wordkstation is not synchron. Cure: Restart digitizer and ErrorViewer on NX workstation. If the problem occures again reinstall SW.
26118 | ERROR_SENDING_MESSAGE | Error sending a message to NX workstation | Meaning: An error message could not be sent to ErrorViewer on NX wrokstation. Reason: 1. The NX workstation has actually been removed from the network. 2. The NX workstation has actually been shut dwon or user has logged out from NX workstation 3. The ErrorViewer has actually been shut down on NX workstation. 4. The ErrorViewer on NX workstation is started but it has a problem and cannot connect and receive messages 6. The digitizer configuration has actually been modified over WEB insterface , the digitizer now is configured for another NX workstation than the one actually used. Cure: 1. Check network cabling connection to NX workstation and digitizer. 2. Start NX workstation and login. 3. Start the ErrorViewer on NX workstation. 4. Exit the ErrorViewer and restart it. 5. Install an ErrorViewer with version 2.0 or higher and start the ErrorViewer on NX workstation. 5. Check configuration of NX workstation and digitizer.
26119 | PING_FAILED | Could not connect to NX workstation. | Meaning: Ping of NX workstation failed Reason: 1. The NX workstation is not available in the network. 2. The NX workstation is not started. Cure: 1. Check network cabling connection to NX workstation and digitizer. 2. Start NX workstation and login.
27136 | COULD_NOT_WRITE_COUNTERS_TO_PMM | Error writing counters to PMM flash | Error writing counters to PMM flash
27137 | COULD_NOT_WRITE_COUNTERS_TO_FILE | Error writing counters to file | Error writing counters to file
27138 | COUNTERS_UNDEFINED_DESTINATION | Undefined destiantion for counters | Undefined destiantion for counters
27139 | COULD_NOT_OPEN_COUNTER_FILE | Could not open counter file for writeing | Could not open counter file for writeing
27649 | TIME_DECRYPTION_READ_ERR | Reading the encrypted time failed. | Meaning: Reading the encrypted time for the periodic license failed. Reason: Not all bytes were read from the file which holds the encrypted time. File or file system is corrupted. Cure: Update the periodic license, do a formatting SW installation and set the digitizer time.
27650 | TIME_DECRYPTION_NO_FILE | File which holds the encrypted time is not found. | Meaning: The file which holds the encrypted time for the periodic license is not found. Reason: File was deleted (e.g. by formatting SW installation) or a PMI board is installed, which is not prepared for a digitizer with periodic licenses. Cure: Update the periodic license and set the digitizer time.
27651 | TIME_DECRYPTION_WRONG_TIME | The decryted time is not correct. | Meaning: The decryted time is not correct. Reason: The file which holds the encrypted time was manipulated or a PMI board from another digitizer with a periodic license was installed. Cure: Do a formatting SW installation, update the periodic license and set the digitizer time.
27652 | TIME_DECRYPTION_LIB_ERR | The decryption of the time failed. | Meaning: The decryption of the time for periodic license failed. Reason: SW error. Cure: Install the newest SW version.
27653 | TIME_ENCRYPTION_NO_FILE | File which holds the encrypted time is not found. | Meaning: The file which holds the encrypted time for the periodic license is not found. Reason: File was deleted (e.g. by formatting SW installation) or a PMI board is installed, which is not prepared for a digitizer with periodic licenses. Cure: Update the periodic license and set the digitizer time.
27654 | TIME_ENCRYPTION_OPEN_ERR | File which holds the encrypted time cannot be opened. | Meaning: The file which holds the encrypted time for the periodic license cannot be opened. Reason: File or file system is corrupted. Cure: Do a formatting SW installation, update the periodic license and set the digitizer time correctly.
27655 | TIME_ENCRYPTION_WRITE_ERR | File which holds the encrypted time cannot be written. | Meaning: The file which holds the encrypted time for the periodic license cannot be written. Reason: File or file system is corrupted. Cure: Do a formatting SW installation, update the periodic license and set the digitizer time correctly.
27656 | TIME_ENCRYPTION_LIB_ERR | The encryption of the time failed. | Meaning: The encryption of the time for periodic license failed. Reason: SW error. Cure: Install the newest SW version.
27657 | NO_ACCESS_TO_TIME_ON_DONGLE | No access to time on USB license drive. | Meaning: Cannot access the time on the USB license drive. Reason: No USB license drive available or USB license drive is defective. Cure: Insert USB license drive, replace USB license drive.
27658 | NO_ACCESS_TO_BOXTIME_ON_DONGLE | No access to box time on USB license drive. | Meaning: Cannot access the box time on the USB license drive. Reason: No USB license drive available or USB license drive is defective. Cure: Insert USB license drive, replace USB license drive.
27659 | MISMATCH_DIGITIZER_DONGLE_TIME | Mismatch between digitizer time and time on USB license drive. | Meaning: Mismatch between digitizer time and time on USB license drive. Reason: The digitizer time is invalid. Cure: Update the periodic license and set the digitizer time correctly.
27660 | NO_DONGLE_FOUND | No USB license drive found. | Meaning: The USB license drive was not found. Reason: No USB license drive available or USB license drive is defective or a standard USB flash memory without license is inserted. Cure: Insert USB license drive, replace USB license drive.
27661 | NO_AGFA_LICENSE_ON_DONGLE | No Agfa license on USB license drive found. | Meaning: The USB license drive does not contain an Agfa license. Reason: The USB license drive does not contain an Agfa license. Cure: Check license information on USB license drive and create a correct periodic license.
27662 | NO_ARCUS_LICENSE_ON_DONGLE | No digitizer license on USB license drive found. | Meaning: The USB license drive does not contain an Agfa digitizer license. Reason: The USB license drive does not contain an Agfa digitizer license. Cure: Check license information on USB license drive and create a correct periodic license.
27663 | WRONG_SERIAL_NUMBER | The digitizer serial number on USB license drive is wrong. | Meaning: The digitizer serial number on the USB license drive is wrong. Reason: Wrong USB license drive inserted. Cure: Check license information on USB license drive and create a correct periodic license.
27664 | LICENSE_EXPIRED | The license on USB license drive has expired. | Meaning: The periodic license on the USB license drive has expired. Reason: Periodic license expired. Cure: Update the periodic license and set the digitizer time correctly.
27665 | LICENSE_CHECK_LIB_ERR | The license check failed. | Meaning: The check of the periodic license failed. Reason: SW error. Cure: Install the newest SW version.
27666 | UPDATE_SYSTEM_TIME_FAILED | Updating the system time of USB license drive failed. | Meaning: Updating the system time of the USB license drive failed. Reason: HW error. Cure: Replace the USB license drive.
27667 | OPEN_HANDLES | Open handles found. | Meaning: Not all handles were released in the Wibu library. Reason: SW error. Cure: Install the newest SW version.
27668 | DIGITIZER_TIME_MANIPULATED | The digitizer time was manipulated. | Meaning: The digitizer time was manipulated. Reason: The digitizer time is invalid. Cure: Update the periodic license and set the digitizer time correctly.
27669 | TOO_MANY_BOX_ENTRIES | Too many entries found on USB license drive. | Meaning: The USB license drive contains too many entries. Reason: Too many entries were stored on the the USB license drive. Cure: Check the entries on the USB license drive and remove unecessary entries. Replace the USB license drive.
27670 | TOO_MANY_INFO_ENTRIES | Too many entries found on USB license drive. | Meaning: The USB license drive contains too many entries. Reason: Too many entries were stored on the the USB license drive. Cure: Check the entries on the USB license drive and remove unecessary entries. Replace the USB license drive.
27671 | DECRYPT_TIME_NULLPTR | Time decryption failed due to a SW error. | Meaning: A null pointer was passed to the SW method. Reason: SW error. Cure. Install the newest SW version.
27672 | GET_DONGLE_TIMES_NULLPTR | Getting dongle times failed due to a SW error. | Meaning: A null pointer was passed to the SW method. Reason: SW error. Cure. Install the newest SW version.
ARC1200 | ARC1200 | -- | IP-transport error. Meaning: Movement of the Handling has not finished successfully. The control reached an unexpected situation. Check the description of the Error Code for more specific information. Reason 1: Error during operation. Cure 1: Restart digitizer to perform the startup selftest procedure. Watch out for messages, e.g. if an IP has to be removed. Reason 2: Software problem. Cure 2: Load new Software. Reason 3: 'Power Board' defective. Cure 3: Check 'Power Board' and replace if defective.
ARC1201 | ARC1201 | -- | No IP present at IP Sensor at the end of the tray. Meaning: Transport of the IP to the 'IP Detection Light Sensor' not successful. At the IP Detection Light Sensor no IP is detected after the scan movement. Reason 1: Transport of the IP not successful. Cure 1: Restart Digitizer. Reason 2: IP has jammed. Cure 2: Open Digitizer and remove jammed IP.
ARC1202 | ARC1202 | -- | Unexpected IP found at the beginning of a new cycle. IP present at IP Detection Light Sensor. Meaning: At the last cycle the expected transport of the IP into the Cassette was not successful. At the 'IP Detection Light Sensor' an IP is detected at the start movement of the new cycle. Reason 1: Handling of the IP not successful. Cure 1: Restart Digitizer. Reason 2: IP has jammed. Cure 2: Open Digitizer and remove jammed IP.
ARC1203 | ARC1203 | -- | Clean Up procedure failed. Meaning: The IP is not transported into the cassette successfully. Reason 1: Handling of the IP not successful. Cure 1: Restart Digitizer. Reason 2: Cassette not inserted straight into 'Cassette Unit'. The IP has been lost at Slow Scan Drum and is placed at IP-Tray. Cure 2: Check the Cassette for straight entry. The distances of the arrow marks at tube side of the inserted Cassette have to be equal to the slot. In case of repeated appearance of this failure the 'Cassette Unit' has to be replaced. Open the Digitizer and remove the lost IP. Reason 3: IP has jammed. Cure 3: Open Digitizer and remove jammed IP. Reason 4: The IP is not transported straight through the Digitizer. The IP has been lost at Slow Scan Drum and is placed at IP-Tray. Cure 4: Open the Digitizer and remove the lost IP. Check by help of Diagnose Cycle the 'Cassette Unit' for feeding the IP straight into the Digitizer. In case of repeated appearance of this failure the 'Cassette Unit' has to be replaced.
ARC1204 | ARC1204 | -- | Return IP procedure failed. Meaning: At a handling error situation the IP is not transported into the cassette successfully. Reason 1: Handling of the IP not successful. Cure 1: Restart Digitizer. Reason 2: Cassette not inserted straight into 'Cassette Unit'. The IP has been lost at Slow Scan Drum and is placed at IP-Tray. Cure 2: Check the Cassette for straight entry. The distances of the arrow marks at tube side of the inserted Cassette have to be equal to the slot. In case of repeated appearance of this failure the 'Cassette Unit' has to be replaced. Open the Digitizer and remove the lost IP. Reason 3: IP has jammed. Cure 3: Open Digitizer and remove jammed IP. Reason 4: The IP is not transported straight through the Digitizer. The IP has been lost at Slow Scan Drum and is placed at IP-Tray. Cure 4: Open the Digitizer and remove the lost IP. Check by help of Diagnose Cycle the 'Cassette Unit' for feeding the IP straight into the Digitizer. In case of repeated appearance of this failure the 'Cassette Unit' has to be replaced.
ARC1205 | ARC1205 | -- | IP transport problem. Meaning: IP transport to scan drum failed. Or transport of IP back to cassette not successful. Root Cause stored in Infocounter Error History. Reason 1: Transport of the IP failed. Cure 1: Restart Digitizer. Reason 2: Cassette not inserted straight into 'Cassette Unit'. The IP has been lost at Slow Scan Drum and is placed at IP-Tray. Cure 2: Check the Cassette for straight entry. The distances of the arrow marks at tube side of the inserted Cassette have to be equal to the slot. In case of repeated appearance of this failure the 'Cassette Unit' has to be replaced. Open the Digitizer and remove the lost IP. Reason 3: IP has jammed. Cure 3: Open Digitizer and remove jammed IP. Reason 4: The IP is not transported straight through the Digitizer. The IP has been lost at Slow Scan Drum and is placed at IP-Tray. Cure 4: Open the Digitizer and remove the lost IP. Check by help of Diagnose Cycle the 'Cassette Unit' for feeding the IP straight into the Digitizer. In case of repeated appearance of this failure the 'Cassette Unit' has to be replaced.
ARC1206 | ARC1206 | -- | Transport Motors stop as Cassette Release Button is pressed during IP transport. Meaning: The mechanical Cassette Release Button is pressed while the Digitizer is still working. The Motors stop to avoid any damage to IP, Cassette or Digitizer. Manual intervention needed. Reason 1: Pressing of the cassette release button before end of cycle. Cure 1: Please clamp the cassette again by pushing it towards the digitizer. Restart Digitizer. Reason 2: 'Cassette Detection Light Sensor' does not work properly. The Light Sensor is not interrupted by the corresponding lever during IP transport. Cure 2: Check Cassette Detection Light Barrier and the corresponding lever. Check mounting of the 'Cassette Detection Light Sensor'. Reason 3: 'Cassette Detection Light Sensor' does not work properly. Cure 3: Check cable and connectors to 'Cassette Detection Light Sensor'. Replace 'Y-Cable Power Board to Drawer/Cass. LS'. Reason 4: 'Cassette Detection Light Sensor' defective. Cure 4: Replace 'Cassette Detection Light Sensor'. Reason 5: The 'Cassette Unit' has a mechanical defect at the sensor lever. The Cassette Detection Light Sensor lever is not moving sufficiently towards the sensor gap. Cure 5: Check the 'Cassette Unit' for proper mechanical function when a Cassette is inserted. Replace 'Cassette Unit'. Reason 6: 'Power Board' defective. Cure 6: Replace 'Power Board'.
ARC1300 | ARC1300 | -- | Decoded Barcode data is invalid. Meaning: Decoding of the barcode on the Image Plate failed. The Digitizer is out of operation. Reason: 1) Image Plate does not arrive in the designated time frame at the position 'read barcode label' due to any transport issue of the Image Plate out of the cassette and inside the Digitizer. This can also be caused by defects of the cassette, the Image Plate or Digitizer components. 2) Image plate arrives only partly at the position 'read barcode label', so that the barcode can not be properly decoded. 3) Barcode is detected, but can not be decoded due to ambient light or due to partial shadowing of the barcode e.g. by the cleaning brush. 4) Barcode is detected, but can not be decoded due to any defects of the barcode label like scratches, contamination, inhomogeneous colouring or bubbles. 5) Barcode is detected, but can not be decoded due to improper position of the barcode-label on the Image Plate or due to improper cutting of the barcode-label. 6) Laser beam out of position, mostly in combination with error code 10852. Cure: 1) Install the most recent digitizer software, refer to the relevant software installation instruction. 2) Check if the cleaning brush is in the home position. 3) Check if the Digitizer is not exposed to direct sunlight. 4) Check if the problem is persistent for all used cassettes. If the problem appears only with specific cassettes: 4a) Check the barcode label on the Image Plate for any defects like scratches, contamination, inhomogeneous colouring or bubbles. If required, clean the barcode label. 4b) Check the proper position of the barcode on the Image Plate. 4c) Check the proper position of the barcode on the barcode-label itself, see also SB13 (DD+DIS078.13E). 4d) Check if the Image Plate is flat, especially on the edges (no damage, not curved), see also SB12 (DD+DIS073.13E). 4e) Check the rear of the Image Plate for any contamination. 4f) Replace the affected Image Plate, refer to the Service Manual for order and replacement. 5) Check the proper position of the cassette when clamped, e.g cassette correctly clamped, no load on the clamped cassette, proper position of the front cover. 6) Check the proper transport of the Image Plate inside the Digitizer: 6a) Check, if the Image Plate can be taken out of the cassette smoothly. If the cassette shows any defects, exchange the cassette, refer to the Service Manual for order and replacement. 6b) Check, if the drawer unit is working properly, e.g. damaged or detached magnets, damages of the drawer fingers, stiff movement of the drawer, locking screw on the motor shaft is not properly fixed. If required exchange the DRAWER UNIT, refer to the Service Manual for order and replacement. 6c) Check the transport path of the Image Plate for any obstacles like metallic particles or broken light collector. If the light collector is broken, exchange the PMT WITH LIGHT COLLECTOR, refer to the Service Manual for order and replacement. 7) If problem is still present, create a backup-file and contact your local support organisation. Only replace the optics module on advise of your local support organisation.
ARC1301 | ARC1301 | -- | Detection of a Barcode Label failed. No recognition of a barcode at all. Meaning: Detection of the barcode on the Image Plate failed. The Digitizer is out of operation. Reason: 1) Image Plate does not arrive in the designated time frame at the position 'read barcode label' due to any transport issue of the Image Plate out of the cassette and inside the Digitizer. This can also be caused by defects of the cassette, the Image Plate or Digitizer components. 2) Barcode can not be detected due to ambient light or due to partial shadowing of the barcode e.g. by the cleaning brush. 3) Barcode can not be detected due to any defects of the barcode label like scratches, contamination, inhomogeneous colouring or bubbles. 4) Barcode can not be detected due to improper position of the barcode-label on the Image Plate or due to improper cutting of the barcode-label. 5) Laser beam not present or out of position, mostly in combination with error code 10852. Cure: 1) Install the most recent digitizer software, refer to the relevant software installation instruction. 2) Check if the cleaning brush is in the home position. 3) Check if the Digitizer is not exposed to direct sunlight. 4) Check if the problem is persistent for all used cassettes. If the problem appears only with specific cassettes: 4a) Check the barcode label on the Image Plate for any defects like scratches, contamination, inhomogeneous colouring or bubbles. If required, clean the barcode label. 4b) Check the proper position of the barcode on the Image Plate. 4c) Check the proper position of the barcode on the barcode-label itself, see also SB13 (DD+DIS078.13E). 4d) Check if the Image Plate is flat, especially on the edges (no damage, not curved), see also SB12 (DD+DIS073.13E). 4e) Check the rear of the Image Plate for any contamination. 4f) Replace the affected Image Plate, refer to the Service Manual for order and replacement. 5) Check the proper position of the cassette when clamped, e.g cassette correctly clamped, no load on the clamped cassette, proper position of the front cover. 6) Check the proper transport of the Image Plate inside the Digitizer: 6a) Check, if the Image Plate can be taken out of the cassette smoothly. If the cassette shows any defects, exchange the cassette, refer to the Service Manual for order and replacement. 6b) Check, if the drawer unit is working properly, e.g. damaged or detached magnets, damages of the drawer fingers, stiff movement of the drawer, locking screw on the motor shaft is not properly fixed. If required exchange the DRAWER UNIT, refer to the Service Manual for order and replacement. 6c) Check the transport path of the Image Plate for any obstacles like metallic particles or broken light collector. If the light collector is broken, exchange the PMT WITH LIGHT COLLECTOR, refer to the Service Manual for order and replacement. 7) If problem is still present, create a backup-file and contact your local support organisation. Only replace the optics module on advise of your local support organisation.
ARC1302 | ARC1302 | -- | No Scan Data of the Barcode Label available at all. Meaning: No data present after Barcode Label scanning procedure. Reason: Latest digizizer software not installed Cure: Install the latest digitizer software, refer to the relevant software installation instruction. If the problem remains, perform a complete backup via the Service User Interface and contact your Regional Support Network for further assistance.
ARC2000 | ARC2000 | -- | The preventive maintenance interval has expired. Meaning: The preventive maintenance interval has expired. Reason 1: The configured interval for preventive maintenance has expired. Cure 1: Perform the preventive maintenance and confirm the completion of the maintenance task .
ARC2100 | ARC2100 | -- | 'Optic Module' eLabel data are not accessible. Meaning: The access to data stored on the 'Optic Module' eLabel failed. Reason 1: The cable to 'Optic Module' is defective or disconnected. Cure 1: Check cables and connectors in the wiring between 'PMI Board' and 'Optic Module'. Replace 'Cable PMI Board to Optic Module'. Reason 2: The 'Optic Module' data storage eLabel contains invalid data or no data at all. Cure 2: Replace 'Optic Module'.
ARC2101 | ARC2101 | -- | Optic Module operating conditions are not within specification. Meaning: The 'Optic Module' setting and adjustment are out of specification. Or the Galvo parameters are not adjustable. Or the Laser is not working stable. Or the detection of begin of scan failed. Or an optical line synchronisation signal is missing during scan. Reason 1: The 'Optic Module' synchronisation failed because of mechanical shock. Cure 1: Restart Digitizer to get detailed selftest result. Try to scan an IP again to test the 'Optic Module'. Reason 2: The newest digitizer software is not installed. Cure 2: Install the newest digitizer software. Reason 3: The cable to 'Optic Module' is defective or disconnected. Cure 3: Check cable and connectors between 'PMI Board' and 'Optic Module'. Replace 'Cable PMI Board to Optic Module'. Reason 4: The 'Optic Module' is defective. Cure 4: Replace 'Optic Module'. Reason 5: 'PMI Board' defective. Cure 5: Replace 'PMI Board'.
ARC2102 | ARC2102 | -- | Optics module from a licensed device detected during start-up, which is not usable in this device. Meaning: Optics module from a licensed device detected during start-up, which is not usable in this device. The Digitizer is out of operation and requires service support. Reason: 1) Optics module from a licensed device has been mounted, which is incompatible with this device. Cure: 1) Replace the detected optics module by a compatible spare optics module, refer to the Service Manual for order and replacement .
ARC2200 | ARC2200 | -- | The shading calibration cannot be performed using the stored calibration file. Meaning: The stored shading calibration data is not valid any more. Prepare image processing failed. Reason 1: Previous repair with replacement of 'PMT with Light Collector' or 'Optic Module'. Cure 1: Perform a shading calibration. Reason 2: Shading calibration file is corrupt. Cure 2: Perform a shading calibration.
ARC2201 | ARC2201 | -- | Photomultiplier data is not accessible. Meaning: The access to data stored on the 'PMT with Light Collector' flash memory failed. Reason 1: The cables to 'PMT with Light Collector' are defective or disconnected. Cure 1: Check cables and connectors between 'PMI Board' and 'PMT with Light Collector'. Replace 'Data Cable PMI Board to Tube Board'. Replace 'Signal Cable PMI Board to Tube Board'. Reason 2: The 'PMT with Light Collector' flash memory contains invalid data or is empty. Cure 2: Replace 'PMT with Light Collector'.
ARC2202 | ARC2202 | -- | Photomultiplier operating conditions not adjustable. Meaning: At the 'PMT with Light Collector' the setting failed. Or the PMT supply voltage is out of specification. Or the communication to the PMT is not successful. Reason 1: The cables to 'PMT with Light Collector' are defective or disconnected. Cure 1: Restart Digitizer to get relevant selftest result. Try to scan an IP again to test the PMT. Check cables and connectors between 'PMI Board' and 'PMT with Light Collector'. Replace 'Data Cable PMI Board to Tube Board'. Replace 'Signal Cable PMI Board to Tube Board'. Reason 2: The 'PMT with Light Collector' is defective. Cure 2: Replace 'PMT with Light Collector'. Reason 3: 'PMI Board' defective. Cure 3: Replace 'PMI Board'.
ARC2203 | ARC2203 | -- | PMT with light collector from a licensed device detected during start-up, which is not usable in this device. Meaning: PMT with light collector from a licensed device detected during start-up, which is not usable in this device. The Digitizer is out of operation and requires service support. Reason: 1) PMT with light collector from a licensed device has been mounted, which is incompatible with this device. Cure: 1) Replace the detected PMT with light collector by a compatible spare PMT with light collector, refer to the Service Manual for order and replacement.
ARC2300 | ARC2300 | -- | Replace the 'Cassette Unit' due to end of its life time. Meaning: The maintenance plan requests to replace the 'Cassette Unit'. Reason 1: The maximum life time of the 'Cassette Unit' is reached. Cure 1: Replace the 'Cassette Unit'. Check via 'Maintenance Indicator' (Service Menue) whether other components have also to be replaced next time. Bundle these necessary mainentance activities.
ARC2301 | ARC2301 | -- | Lubrication of the 'Cassette Unit' is necessary. Meaning: The maintenance plan requests the lubrication of the 'Cassette Unit'. Reason 1: The 'Cassette Unit' has to be lubricated on regular base. The maximum defined cycles are reached. Cure 1: Lubrication of the 'Cassette Unit' according maintenance instruction.
ARC2302 | ARC2302 | -- | Replace the 'Cover Cassette Unit' due to end of its life time. Meaning: The maintenance plan requests to replace the 'Cover Cassette Unit'. Reason 1: The maximum life time of the 'Cover Cassette Unit' is reached. Cure 1: Replace the 'Cover Cassette Unit'. Check via 'Maintenance Indicator' (Service Menue) whether other components have also to be replaced next time. Bundle these necessary mainentance activities.
ARC2400 | ARC2400 | -- | Replace the 'IP Drawer Module' due to end of its life time. Meaning: The maintenance plan requests to replace the 'IP Drawer Module'. Reason 1: The maximum life time of the 'IP Drawer Module' is reached. Cure 1: Replace the 'IP Drawer Module'. Check via 'Maintenance Indicator' (Service Menue) whether other components have also to be replaced next time. Bundle these necessary mainentance activities.
ARC2404 | ARC2404 | -- | Motor error at the 'IP Drawer Module'. Meaning: During motor selftest of 'IP Drawer Module' an error occurred. Reason 1: The cables are defective or disconnected. Cure 1: Check cables and connectors between 'Power Board' and motor of 'IP Drawer Module'. Replace 'Y-Cable Power Board to Motors'. Reason 2: The motor of 'IP Drawer Module' is defective. Cure 2: Replace the 'IP Drawer Module' Reason 3: 'Power Board' defective. Cure 3: Replace 'Power Board'.
ARC2407 | ARC2407 | -- | Reference at the 'IP Drawer Module' cannot be detected. Meaning: The motor drive of the 'IP Drawer Module' does not reach reference position controlled by 'IP Drawer Light Sensor'. Reason 1: 'IP Drawer Light Sensor' is out of position. Cure 1: Bring 'IP Drawer Light Sensor' in correct position. Reason 2: The cable is defective or disconnected. Cure 2: Check cables and connectors to 'IP Drawer Light Sensor' of 'IP Drawer Module'. Replace 'Y-Cable Power Board to Drawer/Cass. LS'. Reason 3: Sensor defective. Cure 3: Replace 'IP Drawer Light Sensor'. Reason 4: 'IP Drawer Module' motor drive does not move at all or with difficulty. Cure 4: Check motor drive of 'IP Drawer Module' for proper work and replace if damaged.
ARC2408 | ARC2408 | -- | IP Transport error at position of 'IP Detection Light Sensor'. Meaning: 'IP Detection Light Sensor' has wrong status for object inside. Expected status is sensor-free. Reason 1: 'IP Detection Light Sensor' is covered by something. Cure 1: Check if 'IP Detection Light Sensor' is mounted correctly and not covered by something. Reason 2: The cable is defective or disconnected. Cure 2: Check cables and connectors in the wiring to 'IP Detection Light Sensor'. Replace 'Cable Power Board to IP Detection LS'. Reason 3 'IP Detection Light Sensor' defective. Cure 3: Replace 'IP Detection Light Sensor'. Reason 4: 'Power Board' defective. Cure 4: Check 'Power Board'. Replace if defective.
ARC2409 | ARC2409 | -- | Motor movement error of 'IP Drawer Module'. Meaning: Any standard movement of 'IP Drawer Module' not finished successfully. Reason 1: Software problem. Cure 1: Load new Software. Reason 2: 'Power Board' defective. Cure 2: Check 'Power Board' and replace if defective.
ARC2410 | ARC2410 | -- | IP transport fails reference sensor position. Meaning: Drawer does not reach the 'IP Drawer Light Sensor' position, while IP is transported by 'IP Drawer Module'. Reason 1: Image plate damaged. Cure 1: Check Image plate. Replace IP. Reason 2: Image plate blocked by something. Cure 2: Remove blocked IP. Check if IP transport path is free. Reason 3: 'IP Drawer Module' not working correctly. Cure 3: Restart digitizer to get selftest results. Check 'IP Drawer Module' with cables. Replace defective part 'Y-Cable Power Board to Drawer/Cass. LS', 'IP Drawer Light Sensor', 'IP Drawer Module'. Reason 4: 'Power Board' defective. Cure 4: Check 'Power Board'. Replace 'Power Board'.
ARC2411 | ARC2411 | -- | Selftest error of 'IP Drawer Light Sensor'. Meaning: Selftest detects 'IP Drawer Light Sensor' of 'IP Drawer Module' disconnected. Reason 1: Reference sensor is disconnected. Cure 1: Check cable and connectors to 'IP Drawer Light Sensor'. Reason 2: Sensor defective. Replace 2: Replace 'IP Drawer Light Sensor'. Reason 3: 'Power Board' defective. Cure 3: Replace 'Power Board'.
ARC2412 | ARC2412 | -- | Selftest error 'IP Detection Light Sensor'. Meaning: Selftest detects 'IP Detection Light Sensor' disconnected. Reason 1: 'IP Detection Light Sensor' is disconnected. Cure 1: Check cable and connectors to 'IP Detection Light Sensor'. Reason 2: 'IP Detection Light Sensor' defective. Cure 2: Replace 'IP Detection Light Sensor'. Reason 3: 'Power Board' defective. Cure 3: Replace 'Power Board'.
ARC2413 | ARC2413 | -- | Selftest error of 'Cassette Detection Light Sensor'. Meaning: Selftest detects 'Cassette Detection Light Sensor' disconnected. Reason 1: The cable is defective or disconnected. Cure 1: Check cable and connectors to 'Cassette Detection Light Sensor'. Replace 'Y-Cable Power Board to Drawer/Cass. LS'. Reason 2: 'Cassette Detection Light Sensor' defective. Cure 2: Replace 'Cassette Detection Light Sensor'. Reason 3: 'Power Board' defective. Cure 3: Replace 'Power Board'.
ARC2414 | ARC2414 | -- | Selftest error of 'Cassette Detection Light Sensor' and 'IP Drawer Light Sensor'. Meaning: Selftest detects 'Cassette Detection Light Sensor' and 'IP Drawer Light Sensor' disconnected. Both sensors are connected via one Y-cable. Reason 1: The cable is defective or disconnected. Cure 1: Check cables and connectors to 'Cassette Detection Light Sensor' and 'IP Drawer Light Sensor'. Replace 'Y-Cable Power Board to Drawer/Cass. LS'. Reason 2: 'Cassette Detection Light Sensor' defective. Cure 2: Replace 'Cassette Detection Light Sensor'. Reason 3: 'IP Drawer Light Sensor' defective. Cure 3: Replace 'IP Drawer Light Sensor'. Reason 4: 'Power Board' defective. Cure 4: Replace 'Power Board'.
ARC2500 | ARC2500 | -- | Replace the 'Slow Scan Drum with Pulley' due to end of its life time. Meaning: The maintenance plan requests to replace the 'Slow Scan Drum with Pulley'. Reason 1: The maximum life time of the 'Slow Scan Drum with Pulley' is reached. Cure 1: Replace the 'Slow Scan Drum with Pulley'. Check via 'Maintenance Indicator' (Service Menue) whether other components have also to be replaced next time. Bundle these necessary mainentance activities.
ARC2501 | ARC2501 | -- | Replace the 'Slow Scan Plate with Belt Tensioner' due to end of its life time. Meaning: The maintenance plan requests to replace the 'Slow Scan Plate with Belt Tensioner'. Reason 1: The maximum life time of the 'Slow Scan Plate with Belt Tensioner' is reached. Cure 1: Replace the 'Slow Scan Plate with Belt Tensioner'. Check via 'Maintenance Indicator' (Service Menue) whether other components have also to be replaced next time. Bundle these necessary mainentance activities.
ARC2504 | ARC2504 | -- | Motor error at the Slow Scan Unit. Meaning: Motor selftest of the 'Slow Scan Motor' detects an error. Reason 1: The cables are defective or disconnected. Cure 1: Check cables and connectors between 'Power Board' and 'Slow Scan Motor'. Replace 'Y-Cable Power Board to Motors'. Reason 2: The 'Slow Scan Motor' is defective. Cure 2: Replace the 'Slow Scan Motor'. Reason 3: 'Power Board' defective. Cure 3: Replace 'Power Board'.
ARC2600 | ARC2600 | -- | Erasure Unit reports supply voltage error. Meaning: The fuse for supply voltage to the 'Erasure Unit' is defective. 'Power Board' defective. Reason 1: Defective 'Power Board' caused by short circuit at the 'Cable Power Board to Erasure Unit'. Cure 1: Replace 'Power Board' AND 'Cable Power Board to Erasure Unit' together! Reason 2: Defective 'Power Board' caused by short circuit at the 'Erasure Unit'. Cure 2: Replace 'Power Board' AND 'Erasure Unit' together!
ARC2601 | ARC2601 | -- | Erasure Unit reports LED current error. Meaning: The Erasure Unit could not be switched on successfully. LED currents are not OK. Reason 1: 'Cable Power Board to Erasure Unit' defective or disconnected. Cure 1: Check cables and connectors to the 'Erasure Unit'. Replace 'Cable Power Board to Erasure Unit'. Reason 2: 'Erasure Unit' current path broken or partly disconnected. A LED line is defective. Only one colour 'on' during selftest. Cure 2: Check the short cable and connector to the overtemperature protection at connector P2. Let the 'Erasure Unit' cool down if hot. Replace 'Erasure Unit'. Reason 3: 'Power Board' current path interrupted or electronic driver defect. Cure 3: Replace 'Power Board'.
ARC2602 | ARC2602 | -- | Erasure Unit reports LED at standby current error. Meaning: The 'Erasure Unit' could not be switched off successfully. LED currents are not OK. The LEDs may still be on! Reason 1: The 'Power Board' is defective driver or LED sensoric fails. Cure 1: Replace 'Power Board'.
ARC2603 | ARC2603 | -- | Erase failed. Meaning: The Erasure procedure failed in active operation. Reason 1: 'Erasure Unit' or 'Cable Power Board to Erasure Unit' defective. Cure 1: Restart digitizer to get detailed selftest result.
ARC2700 | ARC2700 | -- | Incorrect Power Board version. Meaning: Wrong hardware version for 'Power Board' detected during start-up. This type of digitizer needs a 'Power Board' with increased power output for the erasure unit. Reason 1: 'Power Board' has been mounted, which is incompatible with this type of digitizer. Cure 1: Replace 'Power Board' with compatible version as described in the spare parts list.
ARC2800 | ARC2800 | -- | System time error detected. Meaning: The realtime clock is not usable. The coin cell battery for the realtime clock at the 'PMI Board' is empty. The system time is needed to generate unique identifiers as required by DICOM. Reason 1: The battery is disconnected or got loose. Cure 1: Check proper mounting of the coin cell battery at place B2 on the 'PMI Board'. Check time setting via service tool at item Digitizer-Settings in the webbrowser. Restart digitizer. Reason 2: Battery is flat. Battery voltage is too low. Cure 2: Replace 'PMI Board'. Check time setting via service tool at item Digitizer-Settings in the webbrowser.
ARC2801 | ARC2801 | -- | Connection between 'PMI Board' and 'PMT with Light Collector' failed. Meaning: At the 'PMT with Light Collector' module the tube electronic board detection failed. Reason 1: Cable defective or disconnected. Cure 1: Check cables and connectors between 'PMT with Light Collector' and 'PMI Board'. Replace 'Data Cable PMI Board to Tube Board'. Replace 'Signal Cable PMI Board to Tube Board'. Reason 2: 'PMT with Light Collector' is defective. Cure 2: Replace 'PMT with Light Collector'. Reason 3: 'PMI Board' defective. Cure 3: Replace 'PMI Board'.
ARC2802 | ARC2802 | -- | Connection between 'PMI Board' and 'Optic Module' failed. Meaning: 'Optic Module' electronic board not detected. Reason 1: Cable defective or disconnected. Cure 1: Check cables and connectors between 'PMI Board' and 'Optic Module'. Replace 'Cable PMI Board to Optic Module'. Reason 2: 'Optic Module' is defective. Cure 2: Replace 'Optic Module'. Reason 3: 'PMI Board' defective. Cure 3: Replace 'PMI Board'.
ARC2804 | ARC2804 | -- | Flash memory error at 'PMI Board'. Meaning: The internal flash memory of the 'PMI Board' is corrupted or full. Reason 1: Due to a software error the flash memory was corrupted or is full. Cure 1: Install the newest software version with formatting of the flash memory. Restore previously saved backup data. Reason 2: Flash memory on 'PMI Board' is defective. Cure 2: Replace 'PMI Board'.
ARC2805 | ARC2805 | -- | Check of supply voltages on 'PMI Board' failed. Meaning: The check of the power supply voltages on the 'PMI Board' failed. Supply voltage for CAN, Optic, PMT or internal of PMI is missing or not in a valid range. Reason 1: Cable defective or disconnected. Cure 1: Check cable and connectors between 'Power Board' and the 'PMI Board'. Replace 'Cable PMI Board to Power Board'. Reason 2: 'Power Board' is defective. Cure 2: Replace 'Power Board'. Reason 3: 'PMI Board' is defective. Cure 3: Replace 'PMI Board'.
ARC3100 | ARC3100 | -- | Light inside the digitizer detected. Meaning: Check for light inside the digitizer failed. It is too bright to adjust the Photomultiplier. The optical image sensor is not ready for scan. Reason 1: Cover not closed. Cure 1: Close the covers. Restart digitizer. Reason 2: Ambient Light is too bright. Cure 2: Protect the digitizer from direct sunlight. Restart digitizer.
ARC3200 | ARC3200 | -- | Serial Number consistency check failed. Meaning: USB stick error during repair action. No valid digitizer serial number found on USB flash drive, PMI board or a mismatch between stored device serial numbers. The User has inserted a USB flash drive from another digitizer AND the PMI board has been exchanged. Reason 1: New or incorrect USB flash drive inserted AND PMI board exchanged. Cure 1: Insert correct USB flash drive. Restart the device and follow the instructions. Reason 2: New or incorrect USB flash drive inserted AND PMI board exchanged. Cure 2: Configure Digitizer according to service manual. Restart the digitizer and follow the instructions on processing station.
ARC3201 | ARC3201 | -- | Serial Number consistency check failed. Meaning: USB stick operating conditions error. No valid device serial number found on USB flash memory AND no valid device serial number found on PMI board. The User has inserted a USB flash memory from another digitizer, the USB flash memory is not readable or the USB flash memory is empty AND the PMI board has been exchanged or data on PMI board have been destroyed. Reason 1: New or incorrect USB flash memory inserted AND PMI board exchanged. Cure 1: Use the correct USB flash memory and press confirm to use configuration data from USB flash drive. Reason 2: New or incorrect USB flash memory inserted AND data on PMI board deleted. Cure 2: If you already use the correct USB flash memory and this message is displayed a second time, reinstall the Digitizer software with option format.
ARC3300 | ARC3300 | -- | Device E-Label data corrupt. Meaning: The Device E-Label is corrupt. The identification and serial number is needed for operation. The Device E-Label is fixed at the mounting rack and cannot be exchanged. Reason 1: The 'Device E-Label' connection is disturbed. Cure 1: Check cables and connectors between 'PMI Board' and 'Device E-Label'. Replace 'Cable Device E-Label to PMI Board'. Reason 2: The device eLabel is empty, was formatted or contains invalid data. Cure 2: Replace Digitizer.
ARC3301 | ARC3301 | -- | Device E-Label not accessible. Meaning: The Device E-Label is not detected. The identification and serial number is needed for operation. The Device E-Label is fixed at the mounting rack and cannot be exchanged. Reason 1: The 'Device E-Label' is not connected. Cure 1: Check cables and connectors between 'PMI Board' and 'Device E-Label'. Replace 'Cable Device E-Label to PMI Board'. Reason 2: The 'PMI Board' is defective. Cure 2: Replace 'PMI Board'. Reason 3: The 'Device E-Label' is defective. Cure 3: Replace Digitizer.
ARC3401 | ARC3401 | -- | Access to system time on USB license drive failed. Meaning: Access to system time on USB license drive failed. The Digitizer is out of operation and requires service support. Reason: 1) Data structure on the USB license drive corrupted. 2) USB license drive defect. Cure: 1. Visit 'licensing.healthcare.agfa.net' and and update the license on the USB license drive. If access to the USB license drive failed, replace the USB license drive, refer to the Service Manual for order and replacement. 2. Set the system date and time of the Digitizer accordingly in the Service User Interface, refer to the Service Manual.
ARC4000 | ARC4000 | -- | Identification data are invalid. Meaning: The identification data received from the processing station are corrupt or not present. No scan possible without this data. Or a configuration problem occured. Reason 1: Wrong digitizer type configured at the processing station software. Mismatch between the settings and the connected digitizer. Cure 1: Check the settings for the digitizer type at the processing station software in the 'Device Configuration'. If you configure the workstation via 'digitizer configuration wizard', use the configuration wizard for exactly the type of the connected digitizer. Use the service documentation for further information. Reason 2: Processing station sends wrong identification data. Cure 2: Restart digitizer and processing station, identify again at the processing station. Reason 3: Processing station hangs. Processing station sends no identification data. Digitizer waits to receive with retries but without answer. Cure 3: Restart digitizer and processing station, identify again at the processing station. Reason 4: Processing station configuration is wrong. Processing station sends no identification data. Digitizer waits to receive with retries but without answer. Cure 4: Check that processing station is correctly configured for this digitizer. Check configuration of examination studies. Reason 5: Ethernet connection failure. Cure 5: Check Ethernet connection. Reason 6: Software error at processing station Cure 6: Update software of processing station. Reason 7: Software error at digitizer Cure 7: Update software of digitizer.
ARC4001 | ARC4001 | -- | Software error. Meaning: A software error occurred. Digitizer cannot proceed the current action. Reason 1: A bug in the software at the current action. Cure 1: Restart digitizer to return IP into cassette and release cassette. Install the newest digitizer software. Reason 2: A bug in the software causing system stop. Restart not successfully. Cure 2: Clear IP jam manually. Install the newest digitizer software.
ARC4003 | ARC4003 | -- | Software error Meaning: A software error is causing system stop. Reason 1: A bug in the software causing system stop. Restart not successfully Cure 1: Clear IP jam manually. Install the newest digitizer software.
ARC4004 | ARC4004 | -- | Shading calibration cannot be executed. Meaning: Image processing software module reports an error during line corrections. The shading calibration data is not used successfully. The image may be partly or completely uncorrected. Reason 1: 'Optic Module' or 'PMT with Light Collector' replaced without subsequent shading calibration. Cure 1: Perform a shading calibration according to the service documentation. Reason 2: Shading calibration file is corrupt. Cure 2: Perform a shading calibration according to the service documentation. Reason 3: The stored shading calibration data is not valid any more. The Calibration line does not cover the complete width of the image to be corrected. Cure 3: Perform a shading calibration with the biggest used cassette format. Refer to the service documentation. Reason 4: Calibration line of this image plate type not available, invalid or could not be loaded. Cure 4: Perform a shading calibration with the image plate type, which has generated the error. Refer to the service documentation. Reason 5: Software Bug. Preparation of the correction line is not successful. The image is uncorrected. Cure 5: Install newest software.
ARC4005 | ARC4005 | -- | Software error in the image processing module Meaning: Image processing software module indicates an error performing calibration line corrections. The image may be partly or completely uncorrected. Reason 1: A bug in the software. Cure 1: Restart digitizer to return IP into cassette and release cassette if possible. Remove jammed IP and cassette manually if restart not successful. Install the newest digitizer software.
ARC4006 | ARC4006 | -- | Patient identification data not received from workstation. Meaning: The identification data received from the processing station are not present after timeout. No scan possible without this data. Reason 1: Processing station sends no data. User did not identify at the processing station in time. Cure 1: Identify at the processing station. Reason 2: Processing station hangs. Processing station sends no identification data. Digitizer waits to receive ID data. Retries over a certain period without answer. Cure 2: Restart digitizer and processing station, identify again at the processing station. Reason 3: Processing station configuration is wrong. Processing station does not send ID data. Digitizer waits to receive ID data. Retries over a certain period without answer. Cure 3: Check that processing station is correctly configured. Check configuration of examination studies. Reason 4: Ethernet connection failure. Cure 4: Check Ethernet connection. Reason 5: Software error at processing station Cure 5: Update software of processing station. Reason 6: Software error at digitizer Cure 6: Update software of digitizer.
ARC4007 | ARC4007 | -- | Scanning unit operating conditions error. Meaning: The scanning process failed. the Slow Scan Motor has already stopped. ICC (Image Capture Control) is still waiting for end/begin of the scan, i.e. the required number of pixels has not been stored yet. Reason 1: The cables to 'PMT with Light Collector' are defective or disconnected. Cure 1: Restart Digitizer to get relevant selftest result. Try to scan an IP again to test the PMT. Check cables and connectors between 'PMI Board' and 'PMT with Light Collector'. Replace 'Data Cable PMI Board to Tube Board'. Replace 'Signal Cable PMI Board to Tube Board'. Reason 2: The cable to 'Optic Module' is defective or disconnected. Cure 2: Check cable and connectors between 'PMI Board' and 'Optic Module'. Replace 'Cable PMI Board to Optic Module'. Reason 3: The 'PMT with Light Collector' is defective. Cure 3: Replace 'PMT with Light Collector'. Reason 4: 'PMI Board' defective. Cure 4: Replace 'PMI Board'. Reason 5: The 'Optic Module' is defective. Cure 5: Replace 'Optic Module'.
ARC4008 | ARC4008 | -- | Error information and storage problem while ErrorViewer is disabled. Meaning: Could not save error information file and ErrorViewer is not accessible. Reason 1: USB flash memory not inserted or has no free memory left. Cure 1: Insert a USB flash memory with 10 kBytes free memory left. Reason 2: USB flash memory is write protected or does not allow write access. Cure 2: Remove write protection or use an USB flash memory with 10 kBytes free memory left.
ARC4009 | ARC4009 | -- | Software version mismatch. Meaning: The software of the Digitizer and the ErrorViewer Remote Display on the Image Processing Station does not match. The ErrorViewer Remote Display is not able to display messages from the digitizer. Reason 1: A not matching version of the ErrorViewer Remote Display is installed on the Image Processing Station. Cure 1 : Please install the ErrorViewer Remote Display SW on Image Processing Station which is released for current version of Digitizer software.
ARC4010 | ARC4010 | -- | Connection of Digitizer to image processing station failed. Meaning: Connection between Image Processing Station and Digitizer failed. The ErrorViewer Remote Display on the Image Processing Station is not able to display messages from the Digitizer. Reason 1: The Image Processing Station is not available in the network or has been removed. Cure 1: Check network cable connection between Image Processing Station and Digitizer. Reason 2: The Image Processing Station is not started, shut down or user is not logged in. Cure 2: Start Image Processing Station and login. Reason 3: The ErrorViewer Remote Display on Image Processing Station is not started or shut down or user has logged out. Cure 3: Start the ErrorViewer Remote Display on image processing station. Reason 4: The ErrorViewer Remote Display is started but connection failed. Cure 4: Exit the ErrorViewer Remote Display and restart it. Reason 5: ErrorViewer Remote Display version on Image Processing Station is not supported. Cure 5: Install an ErrorViewer Remote Display with version 2.0 or higher and start the ErrorViewer on Image Processing Station. Reason 6: The Digitizer is configured for another Image Processing Station. Cure 6: Check configuration of Image Processing Station and Digitizer.
ARC4011 | ARC4011 | -- | Software error. Corrupt flash memory. Meaning: A software error occurred. Stored data in the flash memory is corrupt. Reason 1: A bug in the software. Cure 1: Restart digitizer to return IP into cassette and release cassette if possible. Remove jammed IP and cassette manually if restart not successful. Install the newest digitizer software with option format.
ARC4012 | ARC4012 | -- | Software error. Meaning: A software error occurred. Digitizer cannot proceed the current action. Reason 1: A bug in the software at the current action. Cure 1: Restart digitizer to return IP into cassette and release cassette. Install the newest digitizer software. Reason 2: A bug in the software causing system stop. Restart not successfully. Cure 2: Clear IP jam manually. Install the newest digitizer software.
ARC4100 | ARC4100 | -- | The processing station is not available. Meaning: The image processing station is not available. The processing station is not ready to receive the image. Reason 1: The processing station software is not running properly. Cure 1: Check processing station software. Restart processing station. Reason 2: The Ethernet connection is not working properly. Cure 2: Check Ethernet connection. Reason 3: Wrong digitizer or destination setting. Cure 3: Connect to service interface via browser and check the digitizer and destination settings (network addresses and names, AE titles, ports).
ARC4101 | ARC4101 | -- | The image destination is not known. Meaning: The Digitizer cannot send the data to the image processing application. The configured destination is unknown. Reason 1: Wrong Digitizer or destination setting. Cure 1: Connect to service interface via browser and check the Digitizer and destination settings (network addresses and names, AE titles, ports).
ARC4102 | ARC4102 | -- | SW error at the image processing station. Meaning: Software error in the processing station software. The IP cannot be scanned because the IP status is wrong. Digitizer rejects the scan because the processing station did not set the IP status to 'exposed'. Normally the IP status is set to 'exposed' by pressing the 'ID' button on the processing station. Reason 1: Problem with Digitizer configuration. Cure 1: Check if digitizer is configured correctly on the processing station. Reason 2: Problem with software installation at the image processing station. Cure 2: Check processing station software installation. Reinstall the image processing station software. Use the latest software version.
ARC4103 | ARC4103 | -- | The image destination is invalid. Meaning: The Digitizer cannot send the data to the image processing application. The configured destination unknown by the Digitizer or set to invalid values. Reason 1: Wrong Digitizer or destination setting. Cure 1: Connect to service interface via browser and check the Digitizer and destination settings (network addresses and names, AE titles, ports).
ARC4104 | ARC4104 | -- | Read access of reference time on PMI Board failed. Meaning: Read access of reference time on PMI Board failed. The Digitizer is out of operation and requires service support. Reason: 1) The Digitizer system time is out of sync. 2) The time on the USB license drive is out of sync. 3) Corrupted data on the PMI Board. Cure: 1) Restart the Digitizer and check, if the problem persists. 2) Visit 'licensing.healthcare.agfa.net' and update the reference time of the USB license drive. 3) Set the system date and time of the Digitizer accordingly in the Service User Interface, refer to the Service Manual. 4) Reinstall the Digitizer software with formatting.
ARC4105 | ARC4105 | -- | No reference time on PMI Board found. Meaning: No reference time on PMI Board found. The Digitizer is out of operation and requires service support. Reason: 1) No data (e.g. new spare PMI Board installed) or corrupted data on the PMI Board. Cure: 1) Restart the Digitizer and check, if the problem persists. 2) Visit 'licensing.healthcare.agfa.net' and update the reference time of the USB license drive. 3) Set the system date and time of the Digitizer accordingly in the Service User Interface, refer to the Service Manual
ARC4106 | ARC4106 | -- | Reference time on PMI Board not valid. Meaning: Reference time on PMI Board not valid. The Digitizer is out of operation and requires service support. Reason: 1) Corrupted data on the PMI Board 2) A PMI board from another licensed digitizer was installed. Cure: 1) Restart the Digitizer and check, if the problem persists. 2) Do a formatting software installation, refer to the Service Manual. 3) Visit 'licensing.healthcare.agfa.net' and update the reference time of the USB license drive. 4) Set the system date and time of the Digitizer accordingly in the Service User Interface, refer to the Service Manual.
ARC4107 | ARC4107 | -- | Write access of reference time on PMI Board failed. Meaning: Write access of reference time on PMI Board failed. The Digitizer is out of operation and requires service support. Reason: 1) Data structure on the PMI Board corrupted. 2) PMI Board defect. Cure: 1) Restart the Digitizer and check, if the problem persists. If selftest during boot up points to a defect PMI Board, replace the PMI Board, refer to the Service Manual for order and replacement. 2) Do a formatting software installation, refer to the Service Manual. 3) Visit 'licensing.healthcare.agfa.net' and update the reference time of the USB license drive. 4) Set the system date and time of the Digitizer accordingly in the Service User Interface, refer to the Service Manual.
ARC4108 | ARC4108 | -- | Digitizer system time is invalid. Meaning: Digitizer system time is invalid. The Digitizer is out of operation and requires service support. Reason: 1) The Digitizer system time is out of sync or not valid anymore. Cure: 1) Set the system date and time of the Digitizer accordingly in the Service User Interface, refer to the Service Manual.
ARC4200 | ARC4200 | -- | Calculated Image Plate length reached a warning limit. Meaning: The automatic speed calibration reached a warning limit. There may be distortions on scanned image. Reason 1: Relocation or transportation of the digitizer. This can have a slight influence on transport mechanics. Cure 1: Try to repeat new scan cycles with the same IP. The speed calibration is automatically updated during a scan cycle. Reason 2: Ambient Light influences detection of begin of scan (Barcode). Cure 2: Protect the digitizer from direct sunlight. Reason 3: The replacement of modules leads at first scan cycle to a new correction value. Cure 3: Repeat scan cycles with the same IP. The speed calibration is automatically updated during a scan cycle. Reason 4: IP Barcode Label not at the defined position. Cure 4: Repeat scan cycles with the same IP. If the message only happens at a change between different IPs/Cassettes then the difference between the IPs is too large. Check the IPs for proper mounted Barcode Labels. Replace damaged IP. Reason 5: IP damaged. IP transportation not steady and stable. This causes a large variation in the correction from one scan cycle to the next cycle. Cure 5: Check the image plate for damages. Replace damaged IP. Reason 6: Slow Scan mechanics damaged. IP transportation not steady and stable. This causes a large variation in the correction from one scan cycle to the next cycle. Cure 6: Check belt tension of Slow Scan drum drive. Check the Slow Scan mechanics for damages. Check 'Slow Scan Belt'. Check the magnetic drum for damages. Replace 'Slow Scan Plate with Belt Tensioner'. Replace 'Slow Scan Belt'.
ARC4201 | ARC4201 | -- | Calculated Image Plate length reached an error limit. There may be distortions in the scanned image. Meaning: The scan length changed and is far away from the allowed tolerance range. The automatic speed calibration of the Slow Scan drive transmission reached an error limit. Reason 1: IP damaged. IP transportation not steady and stable. This causes a large variation in the correction result from one scan cycle to the next cycle. Cure 1: Please check the image plate for damages. Replace IP. Reason 2: Slow Scan mechanics damaged. IP transportation not steady and stable. This causes a large variation in the correction result from one scan cycle to the next cycle. Cure 2: Check belt tension of Slow Scan drum drive. Check the Slow Scan mechanics for damages. Check 'Slow Scan Belt'. Check the magnetic drum for damages. Replace 'Slow Scan Plate with Belt Tensioner'. Replace 'Slow Scan Belt'.